## Intensities
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (cps) | Mg 24 (cps) | In 115 (cps) | U 238 (cps) | CeO 156 (cps) | Ce 140 (IS) (cps) | Ce++ 70 (cps) | Bkgd 220 (cps)
1 | Daily Performance Check |  | 5/29/2020 10:21:28 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.018 | C:\NexIONData\Method\Daily Performance.mth | 11259.3023 | 159883.0878 | 217038.8413 | 94298.8977 | 4198.3503 | 179785.7759 | 6398.5662 | 2.2
2 | Daily Performance Check |  | 5/29/2020 10:23:57 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.019 | C:\NexIONData\Method\Daily Performance.mth | 11420.6302 | 165944.8744 | 220027.7618 | 95571.2961 | 4155.938 | 181931.2973 | 6255.1025 | 2.1
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (cps) | B 10 (cps) | Al 27 (cps) | Ti 47 (cps) | V-1 51 (cps) | V 51 (cps) | Sc 45 (IS) (cps) | Cr 52 (cps) | Mn 55 (cps) | Fe-1 57 (cps) | Fe-2 57 (cps) | Co 59 (cps) | Ni 60 (cps) | Cu 65 (cps) | Zn 66 (cps) | Ge 72 (IS) (cps) | As 75 (cps) | As-1 75 (cps) | Se 82 (cps) | Sr 87 (cps) | Mo 98 (cps) | Rh 103 (IS) (cps) | Ag 107 (cps) | Cd 111 (cps) | In 115 (IS) (cps) | Sb 121 (cps) | Sn 118 (cps) | Ba 137 (cps) | Tb 159 (IS) (cps) | Tl 205 (cps) | Pb 208 (cps) | Th 232 (cps) | U 238 (cps) | Na 23 (cps) | Mg 24 (cps) | K 39 (cps) | Ca 44 (cps) | Sc-1 45 (IS) (cps) | Cl 35 (cps)
3 | Rinse |  | 5/29/2020 10:31:04 AM |  | C:\NexIONData\DataSet\052920\Rinse.020 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 26.6667 | 4428.6967 | 12391.5668 | 165.3343 | 3184.3601 | 23112.8115 | 11027822.458 | 70275.3313 | 4350.6629 | 29616.8627 | 29616.8627 | 505.3423 | 376.6717 | 5822.5399 | 4002.5647 | 851495.4669 | 481.9311 | 189.3346 | -759.3029 | 34.0571 | 20.4924 | 2388403.3556 | 278.6694 | 985.3207 | 2844639.4889 | 49.3334 | 268.6692 | 570.0115 | 2311183.9 | 18.6667 | 1670.0371 | 236.002 | 451.3414 | 62545.0748 | 31930.7131 | 2270942.1232 | 33864.7965 | 11027822.458 | 554020.2153
4 | Rinse |  | 5/29/2020 10:37:14 AM |  | C:\NexIONData\DataSet\052920\Rinse.021 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 14.6667 | 4481.3706 | 19764.045 | 184.0012 | 3987.2234 | 22268.8043 | 11499499.5829 | 68995.5529 | 5239.6438 | 28690.1203 | 28690.1203 | 1078.7165 | 526.0103 | 6559.508 | 10887.3262 | 960337.2552 | 526.7884 | 278.0028 | -761.1802 | 123.2474 | 30.8668 | 2694409.8184 | 312.0034 | 884.0805 | 3021355.1816 | 52.0001 | 321.337 | 635.3477 | 2560003.2014 | 27.3334 | 2002.7233 | 118.6672 | 511.3467 | 202302.6821 | 39740.9962 | 2321326.975 | 36943.0899 | 11499499.5829 | 614947.4102
5 | Rinse |  | 5/29/2020 10:41:20 AM |  | C:\NexIONData\DataSet\052920\Rinse.022 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 75.3338 | 13420.8111 | 1153153.8538 | 1375.503 | 8698.6046 | 30270.4976 | 12965607.1346 | 83995.7373 | 15093.6004 | 45068.4117 | 45068.4117 | 2819.3799 | 2176.4046 | 10038.3474 | 42029.8606 | 972392.2054 | 337.6014 | 347.338 | -714.6343 | 3283.878 | 364.4964 | 2683443.334 | 798.0417 | 1080.2971 | 3139027.189 | 283.3405 | 9327.4003 | 3883.3277 | 2753731.3644 | 444.6856 | 5954.3732 | 2152.4616 | 2343.2252 | 4993107.2305 | 347425.1428 | 2733382.6387 | 98934.7989 | 12965607.1346 | 690264.9151
6 | Rinse |  | 5/29/2020 10:44:59 AM |  | C:\NexIONData\DataSet\052920\Rinse.023 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 151.3341 | 43809.3733 | 2897436.97 | 2567.5657 | 10852.1285 | 51120.6705 | 8778500.3844 | 138890.5849 | 24995.9454 | 91796.6762 | 91796.6762 | 3728.4931 | 3131.6774 | 10106.9087 | 348666.982 | 618551.5165 | -1021.1203 | 368.0048 | -334.3508 | 8992.7507 | 821.6035 | 1836977.3886 | 2633.5773 | 1197.3117 | 2227760.9861 | 577.345 | 15868.1721 | 5590.4327 | 1946148.2385 | 920.0299 | 59878.0291 | 4514.7169 | 3243.7105 | 12911357.4735 | 947220.6847 | 3943828.9802 | 251307.8792 | 8778500.3844 | 500549.2791
7 | Standard Blank |  | 5/29/2020 10:48:37 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.024 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 13.3333 | 11978.3565 | 124755.6151 | 448.6739 | 3775.8333 | 41709.1981 | 8395262.3414 | 124162.5894 | 5769.8319 | 75395.7753 | 75395.7753 | 442.6735 | 504.0089 | 5339.6654 | 17102.4417 | 607813.0278 | -1076.1175 | 174.0011 | -357.8119 | 584.8317 | 37.2142 | 1766945.5575 | 742.6861 | 857.9043 | 2148266.0525 | 68.0002 | 426.6731 | 712.0178 | 1860629.0307 | 14 | 1286.0214 | 136.6673 | 120.6672 | 337890.0452 | 124122.4059 | 3639436.4848 | 61374.5292 | 8395262.3414 | 440240.9865
8 | Standard Blank |  | 5/29/2020 10:52:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.025 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 19.3333 | 12588.2138 | 198543.8528 | 626.6804 | 4019.2321 | 42252.0655 | 8391115.6227 | 124756.3811 | 6005.9305 | 75271.7888 | 75271.7888 | 397.3391 | 568.0113 | 5342.3322 | 18161.8196 | 608360.5112 | -1067.8634 | 179.3345 | -336.615 | 799.4055 | 37.7347 | 1768327.2477 | 1724.1053 | 833.8337 | 2140832.383 | 69.3335 | 720.6856 | 793.3554 | 1846823.2267 | 38.6667 | 1874.7143 | 427.3398 | 144.0008 | 479096.239 | 145544.66 | 3616808.4513 | 66610.3106 | 8391115.6227 | 437307.404
9 | 20 ppm Ca |  | 5/29/2020 10:55:53 AM |  | C:\NexIONData\DataSet\052920\20 ppm Ca .026 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 17.3333 | 14845.7101 | 130031.1432 | 485.3416 | 2061.4821 | 43098.1263 | 8523689.3714 | 127546.1843 | 9351.0619 | 285645.5814 | 285645.5814 | 4332.6572 | 17547.451 | 5973.9156 | 9647.3019 | 595456.6339 | -964.4037 | 145.3341 | -347.927 | 6039.8049 | 178.5486 | 1738196.7568 | 1475.4102 | 1699.5479 | 2143348.6999 | 183.3345 | 739.3526 | 1183.3824 | 1867642.9893 | 32 | 2124.7277 | 396.0055 | 228.0018 | 304663.3041 | 193397.055 | 3869753.7459 | 67639957.6422 | 8523689.3714 | 454670.7259
10 | 20 ppm Cl |  | 5/29/2020 10:59:32 AM |  | C:\NexIONData\DataSet\052920\20 ppm Cl.027 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 12.6667 | 15891.5083 | 80498.9696 | 433.34 | 16043.0693 | 39579.0439 | 8655978.1848 | 120849.6775 | 6054.6161 | 70811.0638 | 70811.0638 | 320.0036 | 692.0168 | 5156.9306 | 9073.5532 | 595451.1608 | -1022.6441 | 1150.0464 | -352.1354 | 572.9081 | 31.3693 | 1741404.0867 | 9789.3605 | 1159.9052 | 2185668.4196 | 57.3335 | 480.6748 | 784.0215 | 1886245.4239 | 26 | 1654.7022 | 110.0004 | 40.0001 | S | 141260.3193 | 4291329.2369 | 72619.4739 | 8655978.1848 | 6913847.4118
11 | Rinse |  | 5/29/2020 11:03:12 AM |  | C:\NexIONData\DataSet\052920\Rinse.028 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 68.0002 | 44991.7905 | 1791647.1696 | 1788.1147 | 9633.942 | 47447.5825 | 9040216.5326 | 135035.2061 | 13832.1221 | 83535.5602 | 83535.5602 | 1164.0489 | 2430.8789 | 9024.2128 | 83628.924 | 622553.3926 | -1192.0541 | 326.6704 | -360.1005 | 5415.5535 | 366.6985 | 1864792.9716 | 2746.2679 | 1183.4259 | 2311342.4662 | 228.6686 | 8304.5138 | 3205.7024 | 2023360.2171 | 386.0057 | 15879.5611 | 2622.9122 | 1105.3782 | 5618511.7141 | 575422.2952 | 3964066.7285 | 196373.9102 | 9040216.5326 | 489762.8702
12 | Rinse |  | 5/29/2020 11:06:50 AM |  | C:\NexIONData\DataSet\052920\Rinse.029 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 18.6667 | 41321.4387 | 1033445.4625 | 1294.061 | 6345.4184 | 46097.5956 | 9056482.0973 | 135198.7423 | 8228.3804 | 79625.9684 | 79625.9684 | 524.6764 | 950.6995 | 5624.4418 | 21388.417 | 619668.2712 | -1183.7279 | 228.0018 | -353.2133 | 3550.7909 | 89.6619 | 1867158.4378 | 1016.7029 | 1084.1929 | 2279129.6577 | 71.3335 | 1925.4723 | 1403.4046 | 1996200.0451 | 67.3335 | 3261.4922 | 816.6901 | 259.3361 | 991766.8326 | 379587.6074 | 3722311.5099 | 133262.7095 | 9056482.0973 | 468359.8118
13 | Standard Blank |  | 5/29/2020 11:26:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.030 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 16 | 10341.7458 | 39765.365 | 222.6684 | 2276.6024 | 40075.0303 | 8356347.4478 | 120248.6538 | 5599.7641 | 73831.3857 | 73831.3857 | 234.7631 | 588.0383 | 5217.6197 | 13588.5263 | 604844.8901 | -121.0885 | -8.4272 | 10.3381 | 187.6064 | 32.8398 | 1770242.0711 | 494.0086 | 993.6117 | 2136960.4988 | 61.3335 | 407.3391 | 693.3502 | 1861874.2677 | 20 | 1292.6889 | 61.3335 | 25.3334 | 328975.6745 | 93768.8219 | 3856592.4223 | 55614.0421 | 8356347.4478 | 447648.4882
14 | Standard Blank |  | 5/29/2020 11:29:53 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.031 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 16.6667 | 10948.861 | 67047.9684 | 284.6696 | 2308.2178 | 40712.1566 | 8532025.246 | 121696.1548 | 5735.8182 | 72875.4992 | 72875.4992 | 249.9147 | 575.2524 | 5471.0479 | 14355.2208 | 611904.6154 | -45.9609 | -11.461 | 51.8346 | 311.5048 | 35.2142 | 1795566.7559 | 1071.374 | 955.724 | 2164782.5764 | 71.3335 | 622.6803 | 768.6874 | 1909804.5498 | 31.3334 | 1666.7033 | 172.001 | 61.3335 | 439589.1328 | 102547.513 | 3892151.6038 | 58739.8502 | 8532025.246 | 458993.7687
15 | std blank |  | 5/29/2020 11:33:31 AM |  | C:\NexIONData\DataSet\052920\std blank.032 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 16.6667 | 11341.1673 | 71596.8077 | 277.3361 | 2219.1506 | 40918.6583 | 8550612.7306 | 122245.4865 | 5746.4891 | 75378.5198 | 75378.5198 | 291.941 | 582.6963 | 5519.066 | 13045.3015 | 616015.3166 | -196.7752 | -10.6774 | 16.0465 | 291.9705 | 40.5475 | 1806434.3467 | 1100.0424 | 957.6305 | 2188911.9576 | 68.0002 | 603.3461 | 730.6854 | 1885033.0596 | 25.3334 | 1723.3727 | 180.6678 | 56.0001 | 432610.9989 | 103979.2677 | 3891437.4692 | 58301.4071 | 8550612.7306 | 455248.0675
16 | 1ppb |  | 5/29/2020 11:37:10 AM |  | C:\NexIONData\DataSet\052920\1ppb.033 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 8842.069 | 10754.0535 | 191964.9464 | 8802.0442 | 113497.6615 | 150059.075 | 8528690.8962 | 216520.5205 | 158429.021 | 76052.3801 | 76052.3801 | 123537.4141 | 27642.693 | 34893.9049 | 30755.1068 | 610704.2677 | 4944.8848 | 4767.0549 | 1568.7837 | 14204.5109 | 37287.9731 | 1805577.55 | 58887.5988 | 15347.4712 | 2158693.2237 | 39180.3219 | 41612.5199 | 19515.9916 | 1902796.9318 | 70831.1679 | 89802.1022 | 45918.7015 | 79366.5415 | 456239.3808 | 188880.9564 | 4006413.1396 | 66586.1575 | 8528690.8962 | 453596.8247
17 | 10ppb |  | 5/29/2020 11:40:49 AM |  | C:\NexIONData\DataSet\052920\10ppb.034 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 88090.7727 | 40145.6942 | 1056062.153 | 86324.0407 | 1121187.5788 | 1141195.6477 | 8551485.724 | 1092650.999 | 1516641.8557 | 104428.4503 | 104428.4503 | 1227560.471 | 267904.7148 | 280151.1659 | 164915.1512 | 613255.8281 | 52707.8406 | 48828.1595 | 15800.8186 | 140222.9355 | 371072.2894 | 1794386.4281 | 682031.9143 | 147246.228 | 2166355.8936 | 394822.3895 | 424213.7279 | 185171.7419 | 1915316.2118 | 714707.4765 | 890099.3959 | 687029.5225 | 792215.4029 | 1167544.2689 | 765720.4986 | 4955887.5498 | 97900.3526 | 8551485.724 | 449532.1909
18 | 50ppb |  | 5/29/2020 11:44:29 AM |  | C:\NexIONData\DataSet\052920\50ppb.035 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 431942.8972 | 170610.9177 | 4850155.7492 | 429024.0778 | 5651757.4815 | 5579517.0028 | 8834226.4635 | 4890104.4951 | 7549718.8212 | 236800.6851 | 236800.6851 | 6035044.1222 | 1334881.9131 | 1354989.6541 | 792230.1132 | 628346.708 | 252530.3466 | 237378.1562 | 77706.5973 | 702603.1913 | 1823528.0953 | 1846793.8223 | 3358289.884 | 727655.6186 | 2279137.4671 | 1914315.5078 | 2092663.1167 | 919702.5307 | 1951363.1683 | 3467543.4871 | 4410507.14 | 3492544.7694 | 3822203.9048 | 5656489.789 | 3615239.6863 | 10026862.9069 | 310419.9817 | 8834226.4635 | 457256.6819
19 | 100ppb |  | 5/29/2020 11:48:09 AM |  | C:\NexIONData\DataSet\052920\100ppb.036 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 826335.4651 | 325610.4831 | 9621766.2163 | 841152.7158 | 11349312.128 | 11284915.917 | 9050456.0178 | 9668594.9885 | 14987363.0897 | 394177.088 | 394177.088 | 12099756.2501 | 2649119.4022 | 2654501.6278 | 1512021.6228 | 632219.8195 | 481916.583 | 453527.6539 | 151175.559 | 1374873.2549 | 3899075.2857 | 1842753.1572 | 6754493.6082 | 1419238.7666 | 2310382.9939 | 3919052.5375 | 4216430.1148 | 1815614.253 | 2017690.9823 | 7074195.8616 | 8827029.6488 | 7062419.6626 | 7689039.351 | 10972158.0443 | 6879841.8519 | 16286040.198 | 549974.9328 | 9050456.0178 | 448334.0969
20 | 100 PPB |  | 5/29/2020 11:51:50 AM |  | C:\NexIONData\DataSet\052920\100 PPB.037 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 750.0204 | 8079.6177 | 88495.5005 | 1022.037 | 11643.0615 | 47747.1086 | 8565671.8668 | 124165.3065 | 17436.7554 | 384808.4416 | 384808.4416 | 10656.4 | 2715.3234 | 9182.9513 | 7694.7455 | 592505.5605 | 401.5454 | 424.2846 | 191.6172 | 1471.9531 | 3967.7705 | 1746770.3925 | 18493.4678 | 2235.6833 | 2208879.0722 | 3451.755 | 4902.1825 | 2084.1545 | 1874300.4641 | 7302.5502 | 8876.4201 | 7021.7936 | 6486.8366 | 10947258.3011 | 6860707.5741 | 16851988.9169 | 383203.5722 | 8565671.8668 | 432576.3222
21 | 1000 PPB |  | 5/29/2020 11:55:30 AM |  | C:\NexIONData\DataSet\052920\1000 PPB.038 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 84.6669 | 7703.411 | 93632.6828 | 575.345 | 1392.5473 | 40940.6361 | 9008176.4621 | 123796.7724 | 11756.1692 | 3354851.642 | 3354851.642 | 4426.1224 | 1386.3832 | 6356.0812 | 6924.3708 | 616752.7808 | -135.1624 | 19.2588 | 48.7877 | 933.0454 | 593.4548 | 1804949.534 | 4924.1825 | 1287.2745 | 2290628.055 | 272.6693 | 1050.7056 | 813.3565 | 1927800.283 | 600.0126 | 1756.0407 | 2072.8199 | 424.673 | 113423899.1364 | 69846388.1954 | 130603237.2368 | 3708322.1173 | 9008176.4621 | 449744.0506
22 | 5000 PPB |  | 5/29/2020 11:59:08 AM |  | C:\NexIONData\DataSet\052920\5000 PPB.039 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 46.6667 | 7235.834 | 64079.6378 | 1606.7572 | 671.3133 | 39284.8807 | 9449511.0038 | 133700.7536 | 37680.9683 | 16590303.6437 | 16590303.6437 | 17774.8823 | 3959.4598 | 27724.2141 | 6517.4927 | 621262.9748 | -144.538 | 48.7917 | 56.0069 | 3271.7107 | 875.2411 | 1774452.6437 | 2860.2866 | 1386.3234 | 2350929.4724 | 244.0021 | 966.6995 | 30906.0673 | 1951931.6477 | 242.0021 | 1803.3765 | 2345.5313 | 164.001 | 583520184.9074 | 361350239.9195 | 664063340.8649 | 18823472.1082 | 9449511.0038 | 464723.4154
23 | 10000 PPB |  | 5/29/2020 12:02:47 PM |  | C:\NexIONData\DataSet\052920\10000 PPB.040 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 44.0001 | 7568.005 | 105540.1868 | 3006.3164 | 906.7837 | 38498.7068 | 9606461.8809 | 140923.0028 | 71036.1933 | 33260839.1725 | 33260839.1725 | 34971.5571 | 5921.3033 | 16643.6929 | 8135.6649 | 623131.5089 | 84.0587 | 97.5354 | 61.8924 | 6306.1796 | 1121.7149 | 1762254.1912 | 1981.4711 | 1432.2901 | 2316002.9694 | 372.6715 | 1137.3787 | 1495.4117 | 1958097.9946 | 230.6685 | 2233.399 | 1816.7872 | 190.0013 | S | 732299897.759 | S | 38592258.1759 | 9606461.8809 | 467178.2683
24 | 20ppb verification std |  | 5/29/2020 12:06:31 PM | Passed | C:\NexIONData\DataSet\052920\20ppb verification std.041 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 172864.5479 | 8251.7187 | 2057097.2108 | 610.6799 | 2394937.924 | 2425741.7064 | 9059313.2634 | 2098647.7489 | 3095103.0221 | 6382183.6551 | 6382183.6551 | 2454993.6551 | 518729.9849 | 537804.9892 | 318231.7086 | 617721.2067 | 103081.1243 | 95560.7485 | 30950.6454 | 29336506.2554 | 740212.083 | 1799199.1599 | 1324203.1141 | 283086.5583 | 2253443.1981 | 777428.1402 | 486.0083 | 367923.1388 | 1941793.6325 | 1379518.4186 | 1727854.3665 | 1375554.8962 | 1528507.7944 | 222677470.4758 | 139585824.2654 | 257680166.3523 | 11797839.5211 | 9059313.2634 | 464480.9283
25 | Ca+B+Sn+Ti check |  | 5/29/2020 12:10:25 PM | Passed | C:\NexIONData\DataSet\052920\Ca+B+Sn+Ti check.042 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 148.0008 | 71030.1661 | 77925.2869 | 173024.2565 | 1307.3273 | 42209.9124 | 8635935.7111 | 122240.1516 | 6997.7163 | 77379.8679 | 77379.8679 | 1722.6673 | 3459.8016 | 5963.2467 | 6052.6186 | 602898.6894 | 20.122 | -7.9883 | 93.9355 | 16606.5601 | 697.715 | 1754671.1998 | 7403.2715 | 1526.612 | 2194116.1217 | 614.6801 | 851053.6942 | 972.0334 | 1863876.9109 | 1134.7117 | 2276.7359 | 2771.6029 | 905.3623 | 527876.4074 | 187784.4959 | 4502482.9026 | 7306160.3601 | 8635935.7111 | 939930.8699
26 | Verification Check |  | 5/29/2020 12:14:49 PM | Passed | C:\NexIONData\DataSet\052920\Verification Check.043 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 173873.7464 | 7820.8172 | 2018296.2775 | 617.3471 | 2372276.1399 | 2366091.9765 | 8998082.4759 | 1961558.4678 | 3082174.0296 | 6424181.6116 | 6424181.6116 | 2458681.5863 | 516638.0555 | 535617.9999 | 317495.7407 | 618327.5819 | 102277.4946 | 95353.4236 | 30872.7883 | 29195664.084 | 730806.6664 | 1762776.0691 | 1311020.4814 | 283186.9621 | 2265578.2667 | 776955.0019 | 1349.3983 | 366597.6025 | 1890310.7419 | 1366168.2074 | 1707705.3906 | 1370585.6141 | 1530261.4916 | 220720472.6359 | 139249695.9205 | 254403150.2256 | 11656588.902 | 8998082.4759 | 462605.4585
27 | Rinse |  | 5/29/2020 12:18:29 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.044 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 154.0009 | 33946.3648 | 614277.9699 | 987.3683 | 2421.95 | 44868.2853 | 9093535.8174 | 130156.2693 | 9611.2715 | 79473.9828 | 79473.9828 | 2089.123 | 1142.7887 | 5703.811 | 12971.4277 | 623832.7021 | -6.5682 | 47.9565 | 74.6793 | 13936.8399 | 991.5701 | 1878759.7737 | 10345.8371 | 1576.9489 | 2338270.6152 | 636.0151 | 2096.8229 | 1426.7402 | 2061052.397 | 1174.0526 | 3797.5401 | 9119.6028 | 1329.4024 | 994906.6353 | 342490.9946 | 4306367.3771 | 132368.692 | 9093535.8174 | 470295.8713
28 | CRM 1_3 |  | 5/29/2020 12:22:07 PM | Passed | C:\NexIONData\DataSet\052920\CRM 1_3.045 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 46042.7719 | 9098.8988 | 3073854.1687 | 509.3425 | 824896.2977 | 859667.4701 | 8750231.1604 | 595052.4982 | 1455378.8219 | 144710.8395 | 144710.8395 | 731248.6303 | 383285.7707 | 137256.9189 | 305976.6912 | 602331.5374 | 100366.9846 | 99071.49 | 4408.6918 | 839383.1118 | 892869.6348 | 1733942.5002 | 29205.9495 | 38319.3164 | 2192922.9262 | 97088.1301 | 641.348 | 218537.7086 | 1821914.7781 | 165559.9099 | 838660.0602 | 7884.8443 | 185738.6011 | 165810471.6118 | 154125162.6422 | 80728702.4422 | 30482221.5623 | 8750231.1604 | 456699.3195
29 | CRM |  | 5/29/2020 12:25:47 PM | Failed | C:\NexIONData\DataSet\052920\CRM.046 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 173950.885 | 15017.3335 | 12717076.9853 | 1067.3734 | 3660374.9661 | 3683711.0568 | 9330874.8373 | 2050058.5835 | 6189444.0212 | 355224.1519 | 355224.1519 | 3069163.1286 | 1472460.9444 | 500659.0435 | 993166.2233 | 593641.0733 | 355391.0072 | 352457.9418 | 14293.3894 | 3781329.2908 | 4082870.9238 | 1726819.3674 | 119509.4846 | 135614.3578 | 2245348.7625 | 369147.6756 | 1293.3922 | 866944.9 | 1904958.8572 | 647003.6607 | 3233749.2783 | 13311.5646 | 754031.7485 | 696081242.9729 | 644957785.6009 | 327843695.9484 | 128029500.7892 | 9330874.8373 | 466609.3581
30 | Rinse |  | 5/29/2020 12:29:27 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.047 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 170.6678 | 40826.389 | 604375.2993 | 918.6972 | 4876.1899 | 49037.0643 | 8995663.5567 | 139316.7602 | 11354.6459 | 81439.41 | 81439.41 | 2690.7545 | 1716.7734 | 5549.7484 | 12887.9083 | 614032.0677 | 47.5699 | 188.8148 | 58.361 | 5566.3869 | 3363.2007 | 1822728.9738 | 5704.491 | 1470.2786 | 2266224.9466 | 506.0101 | 1839.454 | 1602.0929 | 1942846.8254 | 834.027 | 4478.295 | 1040.0387 | 976.04 | 1024641.5977 | 614591.4011 | 4341873.4411 | 175142.3205 | 8995663.5567 | 473479.5027
31 | B, Ca, Ti, Sn check |  | 5/29/2020 12:33:05 PM | Passed | C:\NexIONData\DataSet\052920\B, Ca, Ti, Sn check.048 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 42.0001 | 75306.6583 | 66353.775 | 169319.5378 | 4155.4249 | 40874.641 | 8719191.3948 | 123105.4892 | 5104.9128 | 73379.6285 | 73379.6285 | 396.9911 | 3435.2129 | 5697.8035 | 7031.143 | 602241.7131 | -105.2685 | -35.5464 | 25.537 | 1539.3797 | 313.0805 | 1746967.9545 | 2016.8091 | 1241.76 | 2207609.5997 | 121.3338 | 833279.6867 | 692.0168 | 1874220.5338 | 149.3341 | 1299.3553 | 239.3353 | 107.3337 | 189181.8118 | 97983.7517 | 4148354.7686 | 7183721.0283 | 8719191.3948 | 953290.5495
32 | Rinse |  | 5/29/2020 12:36:45 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.049 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 78.6669 | 38566.7783 | 584829.7465 | 959.3662 | 3895.1778 | 45104.6318 | 8908837.0874 | 133574.9125 | 7644.7293 | 77942.0174 | 77942.0174 | 927.6511 | 901.1522 | 5066.2321 | 15146.8007 | 602653.8566 | -277.332 | 8.8442 | 50.5434 | 3384.0226 | 559.7475 | 1792223.6781 | 1870.7901 | 1263.1811 | 2242276.8135 | 232.0023 | 2860.9566 | 1194.7181 | 1914996.1418 | 317.3377 | 2521.4237 | 1392.7367 | 455.3434 | 633059.6989 | 279891.8128 | 4063526.2018 | 109624.1472 | 8908837.0874 | 467336.6016
33 | 2 ppb verification check |  | 5/29/2020 12:40:23 PM | Passed | C:\NexIONData\DataSet\052920\2 ppb verification check.050 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 18120.1529 | 10245.0071 | 281743.005 | 341.3374 | 225057.906 | 260065.9951 | 8724150.4824 | 312910.5373 | 310288.302 | 683696.4608 | 683696.4608 | 240647.4711 | 53482.4342 | 58572.505 | 37267.8919 | 599043.4688 | 9952.2051 | 9427.1278 | 3024.2213 | 2905086.9364 | 73833.957 | 1757458.0312 | 123471.8273 | 29489.9233 | 2202869.0756 | 77691.36 | 639.3478 | 36968.4456 | 1857566.8202 | 137043.0144 | 174726.186 | 104878.2595 | 153170.8712 | 22225616.8941 | 13974673.0542 | 28692588.4511 | 1047438.1002 | 8724150.4824 | 462774.3163
34 | Rinse |  | 5/29/2020 12:44:04 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.051 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 235.3354 | 38261.9875 | 683293.4166 | 1078.7088 | 5533.8193 | 47954.2619 | 8727042.273 | 135968.6702 | 11455.9877 | 84172.3293 | 84172.3293 | 3345.1393 | 1619.3392 | 6326.0723 | 19120.5425 | 593003.6587 | -34.8141 | 106.8829 | 95.7586 | 10931.5752 | 1700.1446 | 1776719.9033 | 8566.6274 | 1519.6384 | 2224171.8278 | 931.3651 | 3441.7615 | 1740.1093 | 1909150.7079 | 2296.8592 | 6746.6428 | 10381.8219 | 2083.5007 | 1710652.0311 | 480139.3503 | 4459451.8666 | 125065.8771 | 8727042.273 | 464119.5442
35 | Sample blank |  | 5/29/2020 12:47:44 PM | Passed | C:\NexIONData\DataSet\052920\Sample blank.052 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 36.6667 | 9286.3534 | 69738.3607 | 298.0031 | 2298.5471 | 40557.0555 | 8724295.9813 | 121025.1824 | 4816.1458 | 70760.6816 | 70760.6816 | 332.0999 | 900.728 | 6729.5852 | 5817.2002 | 603182.0814 | -70.5098 | -20.2831 | 6.9942 | 506.3211 | 120.0365 | 1759862.103 | 2251.511 | 1126.1559 | 2204403.8618 | 67.3335 | 475.3413 | 704.6841 | 1874428.8613 | 100.667 | 1263.3541 | 317.3369 | 47.3334 | 190130.4994 | 103510.3974 | 4024219.2891 | 55588.6148 | 8724295.9813 | 455582.5635
36 | Rinse |  | 5/29/2020 12:51:25 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.053 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 68.0002 | 36608.2456 | 535092.961 | 819.3574 | 3708.0866 | 45910.6449 | 9007988.5261 | 135042.6155 | 7604.0342 | 77037.6467 | 77037.6467 | 970.4042 | 875.7228 | 5264.9709 | 10470.5776 | 610500.8869 | -351.478 | 9.9729 | 5.7908 | 3600.9208 | 440.7015 | 1820639.9216 | 1972.1374 | 1246.0899 | 2285742.0857 | 217.3354 | 1803.4486 | 1107.3767 | 1969441.5853 | 356.6718 | 2528.0889 | 2444.8767 | 513.3451 | 643711.634 | 271471.9396 | 3996385.7456 | 96729.8443 | 9007988.5261 | 465737.0413
37 | Spike Blank |  | 5/29/2020 12:55:03 PM | Passed | C:\NexIONData\DataSet\052920\Spike Blank.054 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 178111.4758 | 75482.9204 | 2332704.6321 | 173286.0078 | 2402796.8172 | 2413218.7894 | 9151905.1266 | 2047948.6329 | 3125437.8932 | 6679441.4939 | 6679441.4939 | 2518137.7159 | 525732.0223 | 537393.4476 | 324443.5689 | 625353.9207 | 102198.3105 | 95993.1083 | 31080.6142 | 278747.1351 | 740907.9722 | 1810582.184 | 1317005.4878 | 285489.2957 | 2299631.9381 | 777666.8082 | 825382.3367 | 361105.9788 | 1911968.6715 | 1379353.6355 | 1734263.2816 | 1384594.5707 | 1551196.0607 | 234112487.9647 | 144714036.3296 | 262289675.0138 | 7509952.3236 | 9151905.1266 | 470913.855
38 | Rinse |  | 5/29/2020 12:58:43 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.055 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 603.3471 | 37565.4259 | 787085.0438 | 1568.0881 | 9695.8497 | 52189.6879 | 9040962.2145 | 136921.1121 | 20029.0367 | 99804.0902 | 99804.0902 | 9026.4996 | 3307.9375 | 8828.0816 | 35267.9889 | 612748.955 | 162.8316 | 340.237 | 141.7034 | 24686.1199 | 4650.6083 | 1821943.932 | 17268.0989 | 2205.331 | 2298288.8774 | 2659.5944 | 7586.0669 | 3040.9998 | 1978671.2065 | 5728.5407 | 15513.3831 | 21786.3152 | 6026.6725 | 3340993.9843 | 905117.231 | 5240998.411 | 169010.0862 | 9040962.2145 | 477357.6805
39 | Rinse |  | 5/29/2020 1:02:22 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.056 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 84.0003 | 35859.032 | 526788.2673 | 878.0273 | 2334.5959 | 44360.5239 | 8976100.0169 | 130273.3189 | 7722.0925 | 77141.1357 | 77141.1357 | 1064.5913 | 912.5051 | 5168.9359 | 10662.8785 | 606050.694 | -45.3783 | 11.3776 | 74.2264 | 3323.5577 | 486.5573 | 1826448.1193 | 3560.4468 | 1276.6038 | 2308475.0072 | 306.0035 | 1726.7728 | 1113.377 | 1961633.0437 | 458.008 | 2444.7485 | 3913.2039 | 608.0151 | 629267.7288 | 278351.201 | 4102012.7449 | 93626.6276 | 8976100.0169 | 465260.9885
40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99.057 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 31.3334 | 11053.6188 | 164922.2118 | 619.3471 | 1750.2649 | 41265.4829 | 8672274.0617 | 121742.587 | 16650.3658 | 71234.6267 | 71234.6267 | 266.0165 | 440.136 | 8000.2397 | 11686.7822 | 603061.6648 | -265.8366 | -12.097 | 34.3344 | 21779.8509 | 634.6528 | 1734052.0252 | 905.362 | 1069.9455 | 2218558.7952 | 50.0001 | 333.3372 | 12911.1653 | 1858329.3397 | 75.3335 | 1461.3616 | 100.0004 | 2621.5739 | 11116894.7628 | 27051544.5219 | 8940823.534 | 2289267.8845 | 8672274.0617 | 497338.7155
41 | Mi1+HL 200520A 1_99 Duplicate |  | 5/29/2020 1:20:27 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Duplicate.058 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 29.3334 | 12031.0678 | 256521.1372 | 642.0146 | 2276.2175 | 42125.0612 | 8716202.6049 | 124420.809 | 20313.0987 | 75176.6757 | 75176.6757 | 332.2001 | 731.6475 | 8708.6536 | 19041.3517 | 610779.3703 | -187.9421 | -14.8494 | 20.3699 | 22446.7247 | 724.9494 | 1763927.3539 | 2826.2854 | 1120.1706 | 2267236.6017 | 96.667 | 411.3393 | 13181.4129 | 1886952.2236 | 80.0002 | 1766.7083 | 684.0165 | 2803.6085 | 11262469.4278 | 27685006.2218 | 9031512.7262 | 2320493.2937 | 8716202.6049 | 505323.5645
42 | Mi1+HL 200520A 1_9 |  | 5/29/2020 1:24:06 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_9.059 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 37.3334 | 23304.3595 | 297653.8906 | 1561.4189 | 4273.995 | 47545.0528 | 9382278.3866 | 138589.8223 | 130247.9792 | 91028.6063 | 91028.6063 | 229.306 | 2398.5702 | 37541.9454 | 68074.5109 | 623517.7553 | 61.2173 | 5.7901 | 109.5923 | 219296.5928 | 6543.6824 | 1796466.2553 | 4348.6787 | 1160.7632 | 2342338.0792 | 170.6677 | 786.0216 | 116995.1759 | 1962678.5386 | 254.0023 | 5666.4267 | 1626.0929 | 27448.3677 | 116818712.6297 | 286322485.9039 | 56957970.9616 | 23434579.1052 | 9382278.3866 | 968098.7425
43 | Mi1+HL 200520A |  | 5/29/2020 1:27:46 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.060 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 40.0001 | 115958.8928 | 422279.0714 | 9105.5681 | 14699.0437 | 68408.5438 | 9405279.3379 | 192393.7685 | 1171066.6252 | 124322.9486 | 124322.9486 | -3954.3139 | 2742.0964 | 291985.2491 | 547994.0457 | 567308.5209 | 476.0423 | 210.8562 | 336.8707 | 2332791.06 | 63339.7125 | 1593165.5017 | 3783.1826 | 1112.877 | 2114503.6387 | 673.3492 | 2824.9462 | 1077558.3546 | 1813355.8614 | 920.0297 | 38440.0762 | 2176.1658 | 266039.1081 | S | S | 531098248.0392 | 234834480.173 | 9405279.3379 | 6324066.9102
44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.061 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 25.3334 | 114760.4698 | 370262.8451 | 9190.9564 | 17941.8754 | 68177.3661 | 9560951.0371 | 193926.8365 | 1160485.4185 | 128223.1037 | 128223.1037 | -3685.8729 | 4041.3279 | 292084.7973 | 543004.8174 | 571677.5684 | 249.3598 | 226.2715 | 435.221 | 2334671.1607 | 62715.0554 | 1627733.9862 | 2711.5931 | 1101.2474 | 2185164.1999 | 639.3476 | 2685.5858 | 1069299.3977 | 1831224.6104 | 863.3594 | 37864.8588 | 1339.3961 | 259936.9315 | 1142406059.419 | S | 526394745.0998 | 232882662.7634 | 9560951.0371 | 6283493.1852
45 | Rinse |  | 5/29/2020 1:35:06 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.062 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 39.3334 | 33329.5229 | 460096.3341 | 752.0203 | 4491.8197 | 48476.7536 | 8902398.7611 | 144882.3283 | 7775.4578 | 89127.8572 | 89127.8572 | 638.1436 | 804.1359 | 5427.0311 | 8890.1258 | 596122.8232 | -101.6464 | 23.4119 | 77.5286 | 3121.2936 | 296.3317 | 1804895.6194 | 1562.7528 | 1192.5197 | 2238025.3173 | 122.0006 | 1227.3862 | 1496.7482 | 1935787.5391 | 152.6677 | 1867.3803 | 2210.1715 | 442.0081 | 1014314.1518 | 1636365.5066 | 4306627.0624 | 203854.0714 | 8902398.7611 | 473928.6433
46 | Mi1+HL 200520A 1_99 Spike |  | 5/29/2020 1:38:45 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Spike.063 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 164708.7702 | 68763.1468 | 2176202.6112 | 165189.6234 | 2310862.5257 | 2307103.1464 | 8871280.1836 | 1921229.1814 | 2975141.5299 | 6396116.2799 | 6396116.2799 | 2410055.7515 | 509109.6771 | 530055.1974 | 316892.202 | 606982.9094 | 98742.3033 | 92601.8374 | 30015.4914 | 282753.0141 | 706263.8699 | 1740215.2722 | 1264827.703 | 276343.7284 | 2231834.0336 | 752192.7426 | 796521.8924 | 381100.9345 | 1886548.7164 | 1338410.0862 | 1695247.5927 | 1340376.5175 | 1508073.4352 | 231620471.8077 | 162505433.7674 | 257277795.7703 | 9220017.2315 | 8871280.1836 | 508876.9993
47 | Rinse |  | 5/29/2020 1:42:26 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.064 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 172.0011 | 32468.9131 | 583609.9274 | 952.6995 | 4260.5786 | 46757.7254 | 8755917.7424 | 135319.8821 | 9750.0337 | 83722.8874 | 83722.8874 | 2289.6543 | 1200.747 | 5857.8683 | 12072.5326 | 593735.4268 | -137.5519 | 85.3164 | 70.4398 | 4862.1984 | 1223.643 | 1803260.5996 | 9523.2476 | 1368.9872 | 2250819.824 | 692.0178 | 2742.266 | 1648.097 | 1951234.7909 | 1414.0734 | 4133.5733 | 10682.6901 | 1566.0926 | 1114451.1956 | 828186.0408 | 4334555.037 | 135205.0169 | 8755917.7424 | 462707.6269

## Intensities RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 | Mg 24 | In 115 | U 238 | CeO 156 | Ce 140 (IS) | Ce++ 70 | Bkgd 220
1 | Daily Performance Check |  | 5/29/2020 10:21:28 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.018 | C:\NexIONData\Method\Daily Performance.mth | 1.069% | 1.579% | 1.135% | 1.006% | 1.808% | 0.9551% | 1.468% | 31.42%
2 | Daily Performance Check |  | 5/29/2020 10:23:57 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.019 | C:\NexIONData\Method\Daily Performance.mth | 0.8301% | 0.5191% | 0.3535% | 0.6193% | 2.23% | 0.5648% | 0.6449% | 20.7%
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 | B 10 | Al 27 | Ti 47 | V-1 51 | V 51 | Sc 45 (IS) | Cr 52 | Mn 55 | Fe-1 57 | Fe-2 57 | Co 59 | Ni 60 | Cu 65 | Zn 66 | Ge 72 (IS) | As 75 | As-1 75 | Se 82 | Sr 87 | Mo 98 | Rh 103 (IS) | Ag 107 | Cd 111 | In 115 (IS) | Sb 121 | Sn 118 | Ba 137 | Tb 159 (IS) | Tl 205 | Pb 208 | Th 232 | U 238 | Na 23 | Mg 24 | K 39 | Ca 44 | Sc-1 45 (IS) | Cl 35
3 | Rinse |  | 5/29/2020 10:31:04 AM |  | C:\NexIONData\DataSet\052920\Rinse.020 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 38.49% | 15.05% | 23.34% | 12.36% | 14.82% | 5.279% | 14.92% | 4.551% | 3.541% | 9.749% | 9.749% | 9.889% | 15.02% | 16.01% | 10.48% | 15.15% | 47.24% | 15.32% | 5.686% | 28.45% | 32.3% | 18.15% | 12.31% | 18.65% | 9.547% | 14.24% | 17.87% | 10.67% | 13.91% | 16.37% | 9.723% | 4.237% | 45.47% | 7.106% | 13.01% | 4.496% | 3.802% | 14.92% | 13.45%
4 | Rinse |  | 5/29/2020 10:37:14 AM |  | C:\NexIONData\DataSet\052920\Rinse.021 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.83% | 4.886% | 43.44% | 18.09% | 3.199% | 2.173% | 2.648% | 0.8224% | 16.03% | 1.845% | 1.845% | 57.87% | 30.22% | 4.908% | 65.33% | 4.152% | 23.6% | 17.72% | 8.311% | 46.42% | 44.87% | 1.203% | 8.032% | 1.083% | 0.5799% | 36.69% | 14.16% | 15.29% | 0.3442% | 32.99% | 31.22% | 23.25% | 83.09% | 68.55% | 17.54% | 0.6656% | 3.991% | 2.648% | 2.199%
5 | Rinse |  | 5/29/2020 10:41:20 AM |  | C:\NexIONData\DataSet\052920\Rinse.022 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 152.5% | 118% | 171.6% | 153.1% | 85.54% | 43.33% | 11.3% | 30.18% | 125.1% | 61.96% | 61.96% | 153.2% | 147% | 59.15% | 158.4% | 3.466% | 89.16% | 38.84% | 5.159% | 169.9% | 163.6% | 6.94% | 114.2% | 37.06% | 10.33% | 151.8% | 167.5% | 150.8% | 12.19% | 161.5% | 134.3% | 166.4% | 169.1% | 170.9% | 156.4% | 26.74% | 110.9% | 11.3% | 13.19%
6 | Rinse |  | 5/29/2020 10:44:59 AM |  | C:\NexIONData\DataSet\052920\Rinse.023 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 7.63% | 8.23% | 12.87% | 10.38% | 5.453% | 1.593% | 0.8146% | 0.5551% | 8.15% | 1.064% | 1.064% | 14.3% | 6.104% | 2.435% | 14.73% | 0.2978% | 19.21% | 8.98% | 1.811% | 17.86% | 12.81% | 0.8376% | 8.622% | 2.26% | 1.025% | 8.008% | 7.232% | 8.829% | 0.2134% | 11.18% | 18.76% | 8.777% | 19.1% | 5.499% | 13.83% | 0.3377% | 13.12% | 0.8146% | 0.8078%
7 | Standard Blank |  | 5/29/2020 10:48:37 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.024 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 70.89% | 3.184% | 30.75% | 17.15% | 5.487% | 1.047% | 1.596% | 0.7218% | 1.913% | 0.7447% | 0.7447% | 6.49% | 5.339% | 3.816% | 17.58% | 1.844% | 3.348% | 7.178% | 11.72% | 24.75% | 29.89% | 0.8501% | 11.35% | 2.961% | 1.679% | 10.19% | 9.519% | 4.388% | 1.673% | 37.8% | 6.985% | 11.18% | 24.43% | 0.748% | 9.952% | 0.2625% | 5.919% | 1.596% | 1.642%
8 | Standard Blank |  | 5/29/2020 10:52:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.025 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 51.03% | 3.026% | 14.01% | 5.267% | 1.787% | 1.001% | 1.5% | 0.3341% | 4.417% | 0.6382% | 0.6382% | 23.64% | 6.109% | 1.431% | 19.17% | 1.104% | 5.531% | 2.322% | 8.782% | 6.955% | 18.92% | 1.504% | 13.51% | 5.096% | 1.394% | 18.32% | 24.59% | 3.925% | 0.5449% | 38.13% | 22.1% | 8.403% | 34.69% | 21.44% | 4.541% | 0.2951% | 2.11% | 1.5% | 0.863%
9 | 20 ppm Ca |  | 5/29/2020 10:55:53 AM |  | C:\NexIONData\DataSet\052920\20 ppm Ca .026 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 54.53% | 0.6677% | 10.19% | 3.12% | 0.5929% | 0.3799% | 1.189% | 0.1053% | 3.42% | 2.766% | 2.766% | 2.44% | 4.411% | 1.027% | 14.49% | 0.3349% | 24.82% | 26.16% | 6.969% | 0.4094% | 9.95% | 0.8854% | 11.87% | 3.962% | 0.743% | 3.831% | 10.97% | 3.03% | 1.246% | 31.25% | 7.663% | 11.08% | 6.963% | 18.39% | 1.109% | 1.724% | 0.4783% | 1.189% | 0.09892%
10 | 20 ppm Cl |  | 5/29/2020 10:59:32 AM |  | C:\NexIONData\DataSet\052920\20 ppm Cl.027 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 9.116% | 3.615% | 22.32% | 16.53% | 10.5% | 0.8608% | 1.599% | 0.1959% | 0.5312% | 0.2021% | 0.2021% | 5.962% | 8.685% | 0.682% | 5.536% | 1.67% | 23.26% | 6.664% | 9.409% | 12.05% | 42.24% | 2.587% | 5.811% | 6.202% | 2.193% | 17.56% | 2.923% | 4.419% | 1.176% | 13.32% | 6.603% | 17.53% | 25% | S | 5.422% | 0.9234% | 1.642% | 1.599% | 0.969%
11 | Rinse |  | 5/29/2020 11:03:12 AM |  | C:\NexIONData\DataSet\052920\Rinse.028 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 10.6% | 9.062% | 26.59% | 19.51% | 11.36% | 1.071% | 1.249% | 0.1072% | 14.64% | 1.484% | 1.484% | 21.82% | 19.81% | 12.59% | 21.56% | 1.381% | 4.376% | 5.787% | 8.317% | 26.49% | 21.58% | 1.57% | 14.93% | 1.903% | 1.033% | 21% | 25.03% | 19.87% | 0.4334% | 35.8% | 25.54% | 17.2% | 27.24% | 28.51% | 23.68% | 2.699% | 20.96% | 1.249% | 3.26%
12 | Rinse |  | 5/29/2020 11:06:50 AM |  | C:\NexIONData\DataSet\052920\Rinse.029 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 16.37% | 10.42% | 37.17% | 24.91% | 10.03% | 0.5822% | 0.2435% | 0.1451% | 8.492% | 0.7454% | 0.7454% | 14.95% | 24.34% | 4.485% | 40.55% | 1.241% | 10.17% | 10.12% | 11.13% | 39.88% | 18.41% | 1.519% | 3.05% | 4.961% | 0.33% | 16.19% | 32.68% | 22.62% | 1.399% | 38.31% | 47.29% | 7.645% | 48.39% | 39.62% | 30.24% | 0.5516% | 23.91% | 0.2435% | 0.1136%
13 | Standard Blank |  | 5/29/2020 11:26:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.030 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 37.5% | 3.937% | 5.137% | 7.204% | 5.044% | 0.8565% | 1.105% | 0.7208% | 1.213% | 0.6527% | 0.6527% | 8.392% | 1.556% | 2.433% | 12.45% | 1.833% | 158% | 97.77% | 185.3% | 10.33% | 17.01% | 1.8% | 12.35% | 5.853% | 0.5749% | 1.883% | 1.5% | 6.656% | 1.067% | 20% | 7.758% | 13.18% | 16.43% | 2.548% | 2.079% | 0.2384% | 0.1432% | 1.105% | 0.9382%
14 | Standard Blank |  | 5/29/2020 11:29:53 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.031 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 18.33% | 0.9068% | 5.629% | 17.46% | 5.149% | 0.7823% | 1.684% | 0.3793% | 1.915% | 1.689% | 1.689% | 10.66% | 4.857% | 2.473% | 4.969% | 1.162% | 368.7% | 38.67% | 159.3% | 11.1% | 13.33% | 0.9962% | 14.09% | 3.341% | 1.15% | 13.83% | 6.435% | 3.533% | 1.67% | 22.42% | 6.841% | 8.385% | 16.41% | 3.813% | 2.059% | 0.1617% | 0.4651% | 1.684% | 0.3274%
15 | std blank |  | 5/29/2020 11:33:31 AM |  | C:\NexIONData\DataSet\052920\std blank.032 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 24.98% | 1.451% | 6.55% | 13.34% | 2.35% | 0.4907% | 0.975% | 0.5284% | 1.742% | 2.165% | 2.165% | 3.381% | 10.93% | 0.8932% | 6.058% | 2.007% | 25.86% | 22.46% | 34.02% | 3.718% | 14.83% | 2.569% | 5.695% | 5.588% | 1.386% | 8.824% | 4.921% | 2.196% | 0.529% | 61.32% | 4.43% | 17.65% | 12.88% | 2.383% | 3.041% | 0.4774% | 1.415% | 0.975% | 1.252%
16 | 1ppb |  | 5/29/2020 11:37:10 AM |  | C:\NexIONData\DataSet\052920\1ppb.033 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.7377% | 5.181% | 3.643% | 0.4515% | 0.8706% | 0.4874% | 0.2184% | 0.4519% | 0.84% | 0.9373% | 0.9373% | 1.438% | 1.993% | 1.69% | 4.087% | 1.507% | 7.169% | 1.528% | 5.587% | 2.206% | 1.505% | 1.86% | 7.703% | 0.8663% | 1.463% | 0.3177% | 0.6935% | 1.967% | 0.2627% | 0.6739% | 0.624% | 8.513% | 1.139% | 3.112% | 2.466% | 0.2418% | 0.6917% | 0.2184% | 0.4935%
17 | 10ppb |  | 5/29/2020 11:40:49 AM |  | C:\NexIONData\DataSet\052920\10ppb.034 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.6717% | 2.89% | 0.7066% | 0.8324% | 0.6546% | 0.4324% | 4.24% | 0.5672% | 0.578% | 1.525% | 1.525% | 0.9376% | 0.9364% | 1.084% | 0.9146% | 2.742% | 1.447% | 0.5639% | 0.4798% | 1.524% | 1.105% | 4.399% | 1.394% | 1.183% | 5.107% | 0.598% | 0.3088% | 0.6793% | 3.802% | 2.369% | 1.74% | 2.632% | 1.69% | 0.8671% | 1.252% | 0.3126% | 1.367% | 4.24% | 1.473%
18 | 50ppb |  | 5/29/2020 11:44:29 AM |  | C:\NexIONData\DataSet\052920\50ppb.035 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.2934% | 1.653% | 1.439% | 1.164% | 1.458% | 1.59% | 1.49% | 1.347% | 1.384% | 0.9156% | 0.9156% | 1.605% | 0.7497% | 0.7225% | 0.9094% | 0.9125% | 2.124% | 1.5% | 2.307% | 1.559% | 0.2548% | 1.653% | 1.88% | 0.8394% | 0.7325% | 0.8463% | 1.559% | 1.542% | 0.6654% | 2.571% | 1.481% | 1.462% | 2.129% | 0.8369% | 1.509% | 1.295% | 1.129% | 1.49% | 0.2299%
19 | 100ppb |  | 5/29/2020 11:48:09 AM |  | C:\NexIONData\DataSet\052920\100ppb.036 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.349% | 2.521% | 1.035% | 0.8507% | 1.65% | 1.708% | 0.1595% | 1.377% | 2.391% | 1.499% | 1.499% | 1.852% | 1.056% | 1.476% | 0.3927% | 1.426% | 0.4061% | 0.959% | 0.7747% | 2.208% | 2.286% | 0.898% | 0.7489% | 0.9526% | 0.3336% | 1.19% | 0.9164% | 1.25% | 2.162% | 2.487% | 1.583% | 2.019% | 1.92% | 0.627% | 0.8253% | 1.035% | 0.7045% | 0.1595% | 0.8584%
20 | 100 PPB |  | 5/29/2020 11:51:50 AM |  | C:\NexIONData\DataSet\052920\100 PPB.037 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 23.34% | 1.088% | 9.34% | 12.82% | 14.27% | 2.993% | 1.73% | 1.132% | 13.04% | 2.495% | 2.495% | 15.14% | 12.07% | 2.093% | 7.15% | 3.353% | 55.78% | 17.86% | 21.67% | 9.793% | 11.49% | 4.194% | 14.69% | 6.922% | 2.857% | 12.97% | 12.08% | 15.57% | 2.736% | 11.87% | 12.92% | 24.38% | 17.72% | 1.927% | 1.751% | 1.556% | 1.443% | 1.73% | 2.218%
21 | 1000 PPB |  | 5/29/2020 11:55:30 AM |  | C:\NexIONData\DataSet\052920\1000 PPB.038 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 12.12% | 2.987% | 19.25% | 10.72% | 4.052% | 1.134% | 2.596% | 0.7971% | 1.4% | 0.7954% | 0.7954% | 2.343% | 2.646% | 3.046% | 15.37% | 2.528% | 105.2% | 15.83% | 111.8% | 11.74% | 6.212% | 1.762% | 3.259% | 9.866% | 2.63% | 3.388% | 9.931% | 2.829% | 0.935% | 5.033% | 10.84% | 16.99% | 8.121% | 1.487% | 1.234% | 2.026% | 2.384% | 2.596% | 1.917%
22 | 5000 PPB |  | 5/29/2020 11:59:08 AM |  | C:\NexIONData\DataSet\052920\5000 PPB.039 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 32.17% | 4.017% | 12.87% | 5.157% | 13.85% | 0.5241% | 1.331% | 0.5289% | 1.284% | 1.218% | 1.218% | 1.197% | 2.302% | 1.622% | 7.986% | 0.5559% | 89.27% | 13.73% | 34.81% | 4.699% | 5.125% | 0.4984% | 3.582% | 5.571% | 0.8188% | 11.39% | 6.524% | 1.545% | 0.6048% | 7.575% | 5.181% | 20.62% | 21.68% | 1.771% | 1.18% | 1.19% | 2.11% | 1.331% | 0.624%
23 | 10000 PPB |  | 5/29/2020 12:02:47 PM |  | C:\NexIONData\DataSet\052920\10000 PPB.040 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 28.39% | 2.575% | 15.82% | 2.462% | 12.31% | 1.208% | 0.8884% | 0.3531% | 1.186% | 0.3202% | 0.3202% | 1.315% | 2.541% | 2.194% | 10.05% | 0.8836% | 52.09% | 12.84% | 62.14% | 5.598% | 1.44% | 0.2111% | 6.353% | 6.416% | 0.698% | 5.402% | 6.083% | 3.192% | 0.4971% | 8.899% | 2.895% | 25.6% | 25.61% | S | 1.888% | S | 0.9095% | 0.8884% | 0.6754%
24 | 20ppb verification std |  | 5/29/2020 12:06:31 PM | Passed | C:\NexIONData\DataSet\052920\20ppb verification std.041 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 2.078% | 4.241% | 4.348% | 14.44% | 0.5501% | 1.536% | 0.8022% | 5.013% | 1.785% | 1.622% | 1.622% | 1.061% | 0.5454% | 1.135% | 1.052% | 0.8177% | 1.316% | 0.9479% | 1.864% | 1.155% | 0.2371% | 0.802% | 1.776% | 0.7237% | 0.8791% | 1.017% | 12.27% | 0.7704% | 0.9265% | 1.038% | 1.807% | 2.711% | 0.9057% | 1.392% | 1.178% | 1.177% | 0.3127% | 0.8022% | 0.3666%
25 | Ca+B+Sn+Ti check |  | 5/29/2020 12:10:25 PM | Passed | C:\NexIONData\DataSet\052920\Ca+B+Sn+Ti check.042 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 28.38% | 1.275% | 35.12% | 1.134% | 27.87% | 1.963% | 4.39% | 1.246% | 5.027% | 3.523% | 3.523% | 16.31% | 6.881% | 5.304% | 6.215% | 4.693% | 908.7% | 157.6% | 45.84% | 13.73% | 10.92% | 3.701% | 12.62% | 7.86% | 4.098% | 15.97% | 1.747% | 11.41% | 3.564% | 0.833% | 6.914% | 6.231% | 12.5% | 3.85% | 14.81% | 0.4409% | 1.916% | 4.39% | 1.729%
26 | Verification Check |  | 5/29/2020 12:14:49 PM | Passed | C:\NexIONData\DataSet\052920\Verification Check.043 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.5885% | 8.496% | 4.893% | 20.66% | 1.404% | 1.526% | 2.743% | 0.4687% | 1.298% | 0.2956% | 0.2956% | 0.5065% | 0.5981% | 0.8006% | 2.156% | 1.849% | 2.342% | 0.8854% | 2.336% | 1.208% | 0.4048% | 1.144% | 0.9676% | 1.164% | 1.241% | 1.321% | 17.38% | 0.4997% | 2.383% | 1.188% | 0.4586% | 2.018% | 1.118% | 0.7182% | 2.349% | 2.215% | 2.028% | 2.743% | 2.05%
27 | Rinse |  | 5/29/2020 12:18:29 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.044 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 24.68% | 5.473% | 28.16% | 19.47% | 35.79% | 0.8677% | 1.962% | 0.9501% | 13.54% | 0.7091% | 0.7091% | 33.3% | 18.58% | 8.803% | 23.03% | 2.556% | 5007% | 25.22% | 75.25% | 31.07% | 29.5% | 3.612% | 19.23% | 13.74% | 3.885% | 32.32% | 15.13% | 21.95% | 6.275% | 36.69% | 31.94% | 11.62% | 41.74% | 19.79% | 25.66% | 1.173% | 13.11% | 1.962% | 1.644%
28 | CRM 1_3 |  | 5/29/2020 12:22:07 PM | Passed | C:\NexIONData\DataSet\052920\CRM 1_3.045 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 2.336% | 3.298% | 1.627% | 9.733% | 0.8198% | 0.9066% | 1.589% | 0.5993% | 0.2529% | 0.312% | 0.312% | 0.7239% | 0.3327% | 0.7384% | 1.753% | 1.685% | 0.648% | 0.7716% | 1.141% | 0.7707% | 0.3305% | 2.579% | 7.949% | 1.316% | 2.143% | 0.1289% | 15.48% | 1.267% | 0.7494% | 1.156% | 1.284% | 3.904% | 1.864% | 1.423% | 1.992% | 0.9024% | 0.7772% | 1.589% | 1.262%
29 | CRM |  | 5/29/2020 12:25:47 PM | Failed | C:\NexIONData\DataSet\052920\CRM.046 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 1.829% | 14.54% | 1.714% | 7.802% | 1.731% | 2.173% | 1.03% | 6.376% | 1.853% | 2.031% | 2.031% | 1.609% | 1.968% | 2.07% | 1.795% | 1.207% | 1.273% | 1.426% | 0.5674% | 1.623% | 2.098% | 0.9227% | 5.545% | 0.9142% | 0.4134% | 1.136% | 8.971% | 1.066% | 1.956% | 1.186% | 1.166% | 8.831% | 1.385% | 1.776% | 1.918% | 0.9136% | 0.7998% | 1.03% | 0.5086%
30 | Rinse |  | 5/29/2020 12:29:27 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.047 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 43.49% | 5.893% | 33.01% | 22.33% | 31.9% | 3.267% | 2.304% | 1.246% | 21.22% | 2.327% | 2.327% | 43.72% | 31.05% | 7.447% | 15.87% | 2.049% | 619.2% | 53.99% | 123.1% | 45.59% | 37.82% | 1.782% | 15.8% | 6.034% | 1.927% | 43.29% | 16.85% | 22.65% | 2.32% | 40.58% | 34.32% | 17.93% | 54.68% | 29.21% | 38.08% | 2.53% | 25.56% | 2.304% | 0.8333%
31 | B, Ca, Ti, Sn check |  | 5/29/2020 12:33:05 PM | Passed | C:\NexIONData\DataSet\052920\B, Ca, Ti, Sn check.048 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 12.6% | 1.386% | 25.63% | 0.6811% | 5.59% | 1.084% | 0.5617% | 0.4727% | 3.706% | 1.108% | 1.108% | 7.198% | 5.34% | 3.254% | 26.32% | 0.3637% | 74.65% | 23.37% | 296.8% | 10.58% | 11.29% | 0.6342% | 2.999% | 1.312% | 1.203% | 5.789% | 2.11% | 0.8671% | 1.789% | 14.93% | 5.073% | 4.602% | 19.92% | 4.15% | 8.329% | 0.3107% | 0.7972% | 0.5617% | 0.2824%
32 | Rinse |  | 5/29/2020 12:36:45 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.049 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 59.59% | 5.964% | 26.18% | 18.1% | 11.65% | 0.1491% | 0.898% | 0.3861% | 11.4% | 2.175% | 2.175% | 49.22% | 23.24% | 3.186% | 14.24% | 0.9616% | 58.7% | 268.1% | 90.7% | 52.6% | 49.11% | 0.4269% | 10.99% | 12.05% | 0.8623% | 58.64% | 13.44% | 21.1% | 1.452% | 60.42% | 35.39% | 21.71% | 76.08% | 21.96% | 27.79% | 1.172% | 14.48% | 0.898% | 0.5862%
33 | 2 ppb verification check |  | 5/29/2020 12:40:23 PM | Passed | C:\NexIONData\DataSet\052920\2 ppb verification check.050 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 1.464% | 2.459% | 5.019% | 6.099% | 0.6966% | 0.493% | 1.113% | 0.4275% | 0.6026% | 0.5799% | 0.5799% | 1.104% | 2.756% | 1.011% | 1.804% | 0.5122% | 3.41% | 1.004% | 3.283% | 1.762% | 0.883% | 0.8662% | 5.857% | 1.143% | 0.763% | 0.7611% | 12.06% | 1.534% | 0.4294% | 1.912% | 1.431% | 10.29% | 1.388% | 1.997% | 1.06% | 0.7741% | 0.9227% | 1.113% | 0.6247%
34 | Rinse |  | 5/29/2020 12:44:04 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.051 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 33.54% | 6.642% | 31.52% | 23.12% | 22.58% | 1.602% | 2.546% | 0.3904% | 14.33% | 3.025% | 3.025% | 33.65% | 21.5% | 7.498% | 22.26% | 2.472% | 443.1% | 44.79% | 70.2% | 36.82% | 32.46% | 1.034% | 18.66% | 5.579% | 3.053% | 25.93% | 22.22% | 21.7% | 2.302% | 25.39% | 26.6% | 14.23% | 39.08% | 26.7% | 30.01% | 3.06% | 18.36% | 2.546% | 2.423%
35 | Sample blank |  | 5/29/2020 12:47:44 PM | Passed | C:\NexIONData\DataSet\052920\Sample blank.052 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 8.332% | 3.699% | 12.95% | 13.22% | 3.819% | 0.9296% | 0.3505% | 0.94% | 3.756% | 2.527% | 2.527% | 9.566% | 5.453% | 2.267% | 14.26% | 0.1599% | 401.9% | 16.98% | 196.5% | 19.98% | 11.47% | 0.3009% | 4.825% | 5.872% | 1.28% | 7.475% | 12.36% | 8.722% | 1.162% | 16.54% | 6.532% | 8.557% | 8.796% | 3.408% | 5.496% | 0.2633% | 0.8227% | 0.3505% | 0.5472%
36 | Rinse |  | 5/29/2020 12:51:25 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.053 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 48.42% | 4.623% | 23% | 18.38% | 13.49% | 0.7465% | 0.7108% | 0.4075% | 9.025% | 0.9211% | 0.9211% | 43.63% | 8.653% | 3.736% | 17.15% | 0.5303% | 54.3% | 79.12% | 331.6% | 41.39% | 48.42% | 0.8185% | 11.71% | 6.498% | 0.5531% | 59.31% | 13.83% | 11.95% | 0.6643% | 48.24% | 27.01% | 7.962% | 64.17% | 18.97% | 23.21% | 0.4955% | 12.66% | 0.7108% | 0.7326%
37 | Spike Blank |  | 5/29/2020 12:55:03 PM | Passed | C:\NexIONData\DataSet\052920\Spike Blank.054 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.5882% | 1.44% | 2.02% | 0.6473% | 1.446% | 2.647% | 1.701% | 5.231% | 0.4825% | 1.281% | 1.281% | 1.405% | 0.7832% | 0.7318% | 0.4197% | 1.419% | 2.387% | 0.5181% | 1.965% | 0.2694% | 1.011% | 1.442% | 1.586% | 0.2175% | 1.524% | 1.086% | 0.595% | 1.109% | 2.312% | 1.044% | 0.7555% | 1.018% | 1.804% | 1.651% | 1.438% | 1.456% | 0.7929% | 1.701% | 0.477%
38 | Rinse |  | 5/29/2020 12:58:43 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.055 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 34.99% | 5.525% | 21.63% | 18.95% | 26.62% | 4.536% | 0.4772% | 1.386% | 19.05% | 7.08% | 7.08% | 28.27% | 21.94% | 10.85% | 20.16% | 1.132% | 98.83% | 20.99% | 43.96% | 25.08% | 25.86% | 0.9141% | 21.97% | 13.95% | 1.133% | 28.6% | 19.93% | 21.04% | 0.882% | 27.77% | 24.04% | 18.59% | 28.37% | 22.42% | 24.94% | 5.903% | 16.37% | 0.4772% | 1.59%
39 | Rinse |  | 5/29/2020 1:02:22 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.056 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 58.47% | 5.268% | 26.54% | 13.15% | 20.23% | 1.824% | 0.07776% | 0.8025% | 6.57% | 0.2946% | 0.2946% | 38.09% | 12.81% | 3.79% | 29.71% | 0.6161% | 271.7% | 122.9% | 17.91% | 40.55% | 51.89% | 1.099% | 10.37% | 0.6539% | 0.6689% | 34.66% | 16.06% | 10.43% | 1.255% | 37.8% | 25.29% | 6.114% | 50.45% | 17.06% | 24.36% | 0.492% | 14.32% | 0.07776% | 0.1688%
40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99.057 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 25.8% | 6.153% | 14.82% | 18.48% | 16.13% | 1.449% | 1.997% | 1.135% | 1.549% | 2.277% | 2.277% | 4.533% | 12.68% | 1.041% | 3.455% | 2.003% | 32.32% | 51.95% | 29.94% | 0.4593% | 1.831% | 2.252% | 3.432% | 6.563% | 2.296% | 16% | 3.305% | 0.6907% | 1.49% | 20.1% | 5.887% | 10% | 0.552% | 1.259% | 1.931% | 0.4881% | 0.8187% | 1.997% | 1.391%
41 | Mi1+HL 200520A 1_99 Duplicate |  | 5/29/2020 1:20:27 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Duplicate.058 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 10.41% | 3.372% | 4.108% | 13.44% | 4.928% | 0.2572% | 2.475% | 0.4127% | 0.726% | 1.302% | 1.302% | 11.69% | 3.475% | 0.4432% | 2.02% | 2.17% | 79.78% | 32.6% | 308.4% | 0.5213% | 5.378% | 1.781% | 17.78% | 1.114% | 1.17% | 2.389% | 4.87% | 1.693% | 0.4561% | 29.47% | 6.314% | 11.55% | 1.902% | 1.548% | 1.467% | 1.459% | 2.058% | 2.475% | 1.263%
42 | Mi1+HL 200520A 1_9 |  | 5/29/2020 1:24:06 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_9.059 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 29.5% | 5.136% | 18.72% | 6.098% | 6.254% | 1.489% | 2.099% | 1.848% | 2.521% | 2.507% | 2.507% | 21.41% | 3.622% | 2.105% | 1.803% | 0.8347% | 420.3% | 66.99% | 61.45% | 2.267% | 2.882% | 1.41% | 19.6% | 6.29% | 0.9117% | 12.2% | 2.332% | 1.31% | 1.107% | 13.15% | 1.045% | 8.02% | 3.667% | 1.613% | 2.046% | 2.176% | 2.189% | 2.099% | 0.5423%
43 | Mi1+HL 200520A |  | 5/29/2020 1:27:46 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.060 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 21.79% | 2.334% | 11.3% | 1.48% | 7.484% | 1.657% | 2.366% | 1.647% | 0.8967% | 15.96% | 15.96% | 15.06% | 94.44% | 0.9194% | 1.073% | 3.525% | 46.3% | 16.16% | 13.77% | 1.709% | 0.8562% | 1.624% | 21.27% | 5.79% | 6.288% | 1.636% | 3.65% | 0.9953% | 2.307% | 5.112% | 1.741% | 2.946% | 2.194% | S | S | 1.197% | 1.823% | 2.366% | 1.416%
44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.061 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 18.23% | 1.022% | 9.973% | 1.92% | 1.782% | 0.8656% | 0.7956% | 1.133% | 0.6283% | 6.291% | 6.291% | 4.187% | 21.17% | 1.391% | 1.464% | 1.702% | 90.7% | 8.997% | 9.47% | 0.954% | 1.146% | 1.17% | 11.96% | 3.084% | 1.198% | 3.086% | 2.315% | 0.9276% | 1.079% | 3.561% | 0.7067% | 2.553% | 0.6602% | 1.277% | S | 1.057% | 0.7785% | 0.7956% | 0.4874%
45 | Rinse |  | 5/29/2020 1:35:06 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.062 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 62.14% | 2.887% | 24.68% | 20.05% | 6.734% | 2.081% | 1.566% | 0.6531% | 7.978% | 1.545% | 1.545% | 30.31% | 8.607% | 2.63% | 12.14% | 0.9284% | 287.3% | 48.87% | 54.79% | 38.73% | 36.93% | 1.023% | 10.85% | 0.8585% | 0.8173% | 45.9% | 6.257% | 24.36% | 0.9997% | 63.19% | 21.63% | 6.706% | 52.83% | 28.66% | 45.25% | 2.687% | 30.78% | 1.566% | 0.7611%
46 | Mi1+HL 200520A 1_99 Spike |  | 5/29/2020 1:38:45 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Spike.063 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.713% | 2.161% | 0.7528% | 0.8672% | 0.1491% | 0.2449% | 1.542% | 1.04% | 1.762% | 1.947% | 1.947% | 1.211% | 0.6639% | 0.3837% | 0.3849% | 0.4564% | 1.545% | 0.805% | 0.5963% | 1.9% | 1.16% | 1.073% | 1.284% | 0.5641% | 2.221% | 1.457% | 0.6105% | 1.262% | 1.961% | 1.659% | 1.924% | 3.237% | 0.5971% | 0.3776% | 0.4902% | 0.5671% | 1.006% | 1.542% | 0.5079%
47 | Rinse |  | 5/29/2020 1:42:26 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.064 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 32.74% | 4.817% | 30.52% | 22.78% | 22.77% | 1.424% | 0.9838% | 0.2047% | 13.59% | 1.973% | 1.973% | 29% | 11.13% | 3.32% | 17.18% | 0.6667% | 78.14% | 30.34% | 33.43% | 30.64% | 23.09% | 1.527% | 18.76% | 6.533% | 0.4236% | 30.96% | 12.73% | 17.4% | 1.251% | 27.09% | 25.28% | 10.74% | 34.42% | 20.49% | 35.85% | 2.822% | 25.28% | 0.9838% | 0.7558%

## Concentrations
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (mg/L) | Mg 24 (ug/L) | In 115 (ug/L) | U 238 (mg/L) | CeO 156 (mg/L) | Ce 140 (IS) | Ce++ 70 (mg/L) | Bkgd 220 (ug/L)
1 | Daily Performance Check |  | 5/29/2020 10:21:28 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.018 | C:\NexIONData\Method\Daily Performance.mth
2 | Daily Performance Check |  | 5/29/2020 10:23:57 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.019 | C:\NexIONData\Method\Daily Performance.mth
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (ug/L) | B 10 (ug/L) | Al 27 (ug/L) | Ti 47 (ug/L) | V-1 51 (ug/L) | V 51 (ug/L) | Sc 45 (IS) | Cr 52 (ug/L) | Mn 55 (ug/L) | Fe-1 57 (mg/L) | Fe-2 57 (ug/L) | Co 59 (ug/L) | Ni 60 (ug/L) | Cu 65 (ug/L) | Zn 66 (ug/L) | Ge 72 (IS) | As 75 (ug/L) | As-1 75 (ug/L) | Se 82 (ug/L) | Sr 87 (ug/L) | Mo 98 (ug/L) | Rh 103 (IS) | Ag 107 (ug/L) | Cd 111 (ug/L) | In 115 (IS) | Sb 121 (ug/L) | Sn 118 (ug/L) | Ba 137 (ug/L) | Tb 159 (IS) | Tl 205 (ug/L) | Pb 208 (ug/L) | Th 232 (ug/L) | U 238 (ug/L) | Na 23 (ug/L) | Mg 24 (ug/L) | K 39 (ug/L) | Ca 44 (ug/L) | Sc-1 45 (IS) | Cl 35 (ug/L)
3 | Rinse |  | 5/29/2020 10:31:04 AM |  | C:\NexIONData\DataSet\052920\Rinse.020 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
4 | Rinse |  | 5/29/2020 10:37:14 AM |  | C:\NexIONData\DataSet\052920\Rinse.021 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
5 | Rinse |  | 5/29/2020 10:41:20 AM |  | C:\NexIONData\DataSet\052920\Rinse.022 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
6 | Rinse |  | 5/29/2020 10:44:59 AM |  | C:\NexIONData\DataSet\052920\Rinse.023 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
7 | Standard Blank |  | 5/29/2020 10:48:37 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.024 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
8 | Standard Blank |  | 5/29/2020 10:52:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.025 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
9 | 20 ppm Ca |  | 5/29/2020 10:55:53 AM |  | C:\NexIONData\DataSet\052920\20 ppm Ca .026 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
10 | 20 ppm Cl |  | 5/29/2020 10:59:32 AM |  | C:\NexIONData\DataSet\052920\20 ppm Cl.027 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
11 | Rinse |  | 5/29/2020 11:03:12 AM |  | C:\NexIONData\DataSet\052920\Rinse.028 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
12 | Rinse |  | 5/29/2020 11:06:50 AM |  | C:\NexIONData\DataSet\052920\Rinse.029 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
13 | Standard Blank |  | 5/29/2020 11:26:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.030 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
14 | Standard Blank |  | 5/29/2020 11:29:53 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.031 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
15 | std blank |  | 5/29/2020 11:33:31 AM |  | C:\NexIONData\DataSet\052920\std blank.032 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | [6 empty] | 1 | [8 empty] | 1 | [5 empty] | 1 |  |  | 1 |  |  |  | 1 | [8 empty] | 1
16 | 1ppb |  | 5/29/2020 11:37:10 AM |  | C:\NexIONData\DataSet\052920\1ppb.033 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 1 | 1 | 1 | 1 | 1 | 1 | 0.99744 | 1 | 1 | 1 |  | 1 | 1 | 1 | 1 | 0.99138 | 1 | 1 | 1 | 1 | 1 | 0.99953 | 1 | 1 | 0.98619 | 1 | 1 | 1 | 1.0094 | 1 | 1 | 1 | 1 |  |  |  |  | 0.99744
17 | 10ppb |  | 5/29/2020 11:40:49 AM |  | C:\NexIONData\DataSet\052920\10ppb.034 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 9.9996 | 10.1197 | 9.9776 | 10.0007 | 10.0004 | 10.0006 | 1.0001 | 10.0014 | 9.9986 | 10.0544 |  | 9.9992 | 9.9984 | 9.9926 | 9.9825 | 0.99552 | 10.0025 | 10.0018 | 10.0012 | 10.0002 | 10.0003 | 0.99333 | 10.0157 | 10.0014 | 0.9897 | 10.0007 | 10.0029 | 9.9979 | 1.0161 | 10.0003 | 10.0003 | 10.0328 | 9.9993 |  |  |  |  | 1.0001
18 | 50ppb |  | 5/29/2020 11:44:29 AM |  | C:\NexIONData\DataSet\052920\50ppb.035 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 49.8948 | 50.1447 | 49.8679 | 49.9272 | 49.9533 | 49.9467 | 1.0332 | 49.9145 | 49.9483 | 50.121 |  | 49.9175 | 49.9474 | 49.9142 | 49.9968 | 1.02 | 49.859 | 49.8947 | 49.9199 | 49.959 | 49.9069 | 1.0223 | 49.9164 | 49.8814 | 1.0412 | 49.8319 | 49.8709 | 49.8879 | 1.0352 | 49.9023 | 49.9475 | 50.002 | 49.8905 |  |  |  |  | 1.0332
19 | 100ppb |  | 5/29/2020 11:48:09 AM |  | C:\NexIONData\DataSet\052920\100ppb.036 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 98.5071 | 99.2777 | 99.4276 | 99.0529 | 99.5655 | 99.7867 | 1.0585 | 99.8681 | 99.7109 | 99.7337 |  | 99.8954 | 99.6976 | 99.453 | 99.062 | 1.0263 | 98.8216 | 98.8709 | 99.2667 | 99.4109 | 101.3564 | 1.0201 | 100.1313 | 99.1554 | 1.0555 | 100.1319 | 99.8224 | 99.4085 | 1.0704 | 99.69 | 99.3085 | 99.543 | 99.3902 |  |  |  |  | 1.0585
20 | 100 PPB |  | 5/29/2020 11:51:50 AM |  | C:\NexIONData\DataSet\052920\100 PPB.037 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.092644 | -1.0974 | 0.18378 | 0.092845 | 0.087536 | 0.063613 | 1.0018 | 0.074921 | 0.085006 | 104.8994 | 100 | 0.091766 | 0.086912 | 0.15548 | -0.34092 | 0.96184 | 0.13037 | 0.10151 | 0.12296 | 0.092229 | 0.1082 | 0.96697 | 0.27404 | 0.093128 | 1.0091 | 0.090689 | 0.10668 | 0.077583 | 0.99431 | 0.11064 | 0.087043 | 0.10431 | 0.089793 | 100 | 100 | 100 | 100 | 1.0018
21 | 1000 PPB |  | 5/29/2020 11:55:30 AM |  | C:\NexIONData\DataSet\052920\1000 PPB.038 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0080327 | -1.3502 | 0.18766 | 0.033439 | -0.0083206 | -0.019212 | 1.0535 | 0.015512 | 0.040987 | 1058.2292 | 1000.0864 | 0.034995 | 0.031005 | 0.031955 | -0.41667 | 1.0012 | 0.0125 | 0.0067046 | 0.021417 | 0.047406 | 0.014668 | 0.99918 | 0.057938 | 0.019998 | 1.0465 | 0.0052001 | 0.0099861 | 0.0026905 | 1.0227 | 0.0084667 | -8.785e-05 | 0.027814 | 0.0049705 | 1000.2123 | 999.8155 | 999.2398 | 1000.6276 | 1.0535
22 | 5000 PPB |  | 5/29/2020 11:59:08 AM |  | C:\NexIONData\DataSet\052920\5000 PPB.039 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0032111 | -1.6068 | -0.1507 | 0.14666 | -0.014972 | -0.050407 | 1.1051 | 0.11088 | 0.2159 | 5288.8638 | 4999.933 | 0.14682 | 0.12912 | 0.8464 | -0.44637 | 1.0085 | 0.011343 | 0.013206 | 0.026579 | 0.21906 | 0.022553 | 0.9823 | 0.0274 | 0.024562 | 1.074 | 0.0042899 | 0.0074111 | 1.6213 | 1.0355 | 0.0031406 | 0.00021561 | 0.031411 | 0.0014141 | 4996.849 | 4997.4769 | 4998.7367 | 4996.2896 | 1.1051
23 | 10000 PPB |  | 5/29/2020 12:02:47 PM |  | C:\NexIONData\DataSet\052920\10000 PPB.040 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0028459 | -1.5432 | 0.24587 | 0.29911 | -0.013116 | -0.062479 | 1.1235 | 0.18338 | 0.44032 | 10595.7806 | 10003.4904 | 0.29038 | 0.20358 | 0.42124 | -0.33927 | 1.0116 | 0.058892 | 0.023951 | 0.030482 | 0.44109 | 0.029416 | 0.97554 | 0.014084 | 0.029248 | 1.0581 | 0.0076639 | 0.011791 | 0.039482 | 1.0388 | 0.0029667 | 0.0051382 | 0.023674 | 0.0017565 | S | 9992.3616 | S | 10019.3082 | 1.1235
24 | 20ppb verification std |  | 5/29/2020 12:06:31 PM | Passed | C:\NexIONData\DataSet\052920\20ppb verification std.041 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.5855 | -1.1908 | 20.6169 | 0.037318 | 20.9746 | 21.1279 | 1.0595 | 21.1545 | 21.0357 | 2031.1531 | 1917.6143 | 20.7356 | 19.9582 | 20.448 | 20.6405 | 1.0028 | 21.6631 | 21.3196 | 20.787 | 2170.7668 | 19.7069 | 0.99599 | 20.0917 | 20.2229 | 1.0295 | 20.3646 | -0.0032876 | 20.6201 | 1.0301 | 20.1902 | 20.1736 | 20.1351 | 20.5208 | 1986.5761 | 2018.6356 | 2003.7628 | 3236.5155 | 1.0595
25 | Ca+B+Sn+Ti check |  | 5/29/2020 12:10:25 PM | Passed | C:\NexIONData\DataSet\052920\Ca+B+Sn+Ti check.042 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.016314 | 19.7907 | 0.054748 | 21.3486 | -0.0086526 | 0.0085047 | 1.01 | 0.029981 | 0.0096204 | 1.2049 | 1.1375 | 0.012408 | 0.11415 | 0.022066 | -0.46537 | 0.97871 | 0.046883 | 0.00048143 | 0.053774 | 1.2355 | 0.017939 | 0.97135 | 0.099097 | 0.041625 | 1.0024 | 0.014683 | 21.2194 | 0.013705 | 0.98878 | 0.016935 | 0.0069514 | 0.039531 | 0.011865 | 0.85384 | 1.25 | 4.7883 | 2098.2652 | 1.01
26 | Verification Check |  | 5/29/2020 12:14:49 PM | Passed | C:\NexIONData\DataSet\052920\Verification Check.043 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.8554 | -1.3116 | 20.3482 | 0.038325 | 20.9218 | 20.7455 | 1.0523 | 19.6747 | 20.9294 | 2043.205 | 1928.9926 | 20.7501 | 19.8608 | 20.3473 | 20.5691 | 1.0038 | 21.4719 | 21.2548 | 20.7137 | 2158.511 | 19.8591 | 0.97583 | 20.3027 | 20.1202 | 1.035 | 20.2417 | 0.01746 | 20.4361 | 1.0028 | 20.543 | 20.4903 | 20.6135 | 21.1095 | 1983.2846 | 2027.5513 | 1991.8059 | 3219.7821 | 1.0523
27 | Rinse |  | 5/29/2020 12:18:29 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.044 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.016115 | 6.9019 | 5.5637 | 0.081039 | 0.00045382 | 0.012009 | 1.0635 | 0.06824 | 0.025475 | 1.0093 | 0.95286 | 0.014923 | 0.020976 | 0.0041816 | -0.018494 | 1.0127 | 0.038941 | 0.012949 | 0.038302 | 0.99541 | 0.024071 | 1.04 | 0.13331 | 0.038055 | 1.0682 | 0.014122 | 0.033848 | 0.034691 | 1.0934 | 0.015619 | 0.020672 | 0.12294 | 0.015824 | 4.7441 | 3.3309 | 1.3265 | 19.2931 | 1.0635
28 | CRM 1_3 |  | 5/29/2020 12:22:07 PM | Passed | C:\NexIONData\DataSet\052920\CRM 1_3.045 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 5.6746 | -0.82135 | 32.3256 | 0.02744 | 7.4678 | 7.5102 | 1.0233 | 5.2232 | 10.1262 | 23.4619 | 22.1504 | 6.3331 | 15.1205 | 5.1957 | 20.3411 | 0.97779 | 21.6361 | 22.6704 | 3.0283 | 63.6926 | 24.6748 | 0.95987 | 0.44313 | 2.7524 | 1.0018 | 2.6125 | 0.00088994 | 12.5728 | 0.96652 | 2.5821 | 10.4265 | 0.1203 | 2.6568 | 1530.5736 | 2307.6853 | 627.9841 | 8687.4921 | 1.0233
29 | CRM |  | 5/29/2020 12:25:47 PM | Failed | C:\NexIONData\DataSet\052920\CRM.046 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.1146 | 0.81321 | 127.7002 | 0.087407 | 31.1346 | 31.3357 | 1.0913 | 21.5262 | 43.8148 | 94.7094 | 89.4153 | 26.9746 | 58.9908 | 19.8002 | 69.0159 | 0.96368 | 77.6086 | 81.8135 | 9.9844 | 291.1301 | 113.2625 | 0.95593 | 1.8741 | 9.6859 | 1.0258 | 9.7034 | 0.016441 | 48.8205 | 1.0106 | 9.6549 | 38.519 | 0.19583 | 10.3229 | 6038.2467 | 9062.1223 | 2483.0973 | 34263.7294 | 1.0913
30 | Rinse |  | 5/29/2020 12:29:27 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.047 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.018509 | 9.2175 | 5.5777 | 0.074712 | 0.022659 | 0.053854 | 1.052 | 0.1888 | 0.038762 | 2.058 | 1.943 | 0.020523 | 0.044285 | 0.0021365 | -0.0057285 | 0.99678 | 0.052238 | 0.045055 | 0.029223 | 0.39518 | 0.087656 | 1.009 | 0.06901 | 0.034212 | 1.0353 | 0.011405 | 0.029411 | 0.047446 | 1.0307 | 0.011898 | 0.031862 | 0.012541 | 0.012437 | 5.1711 | 7.4185 | 1.9983 | 31.8137 | 1.052
31 | B, Ca, Ti, Sn check |  | 5/29/2020 12:33:05 PM | Passed | C:\NexIONData\DataSet\052920\B, Ca, Ti, Sn check.048 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0030967 | 20.9473 | -0.071951 | 20.6693 | 0.017237 | -0.0078191 | 1.0197 | 0.039493 | -0.0035839 | -0.1036 | -0.097813 | 0.00096688 | 0.11322 | 0.011896 | -0.39686 | 0.97764 | 0.01874 | -0.0057445 | 0.0067846 | 0.095176 | 0.0075056 | 0.96708 | 0.014904 | 0.02021 | 1.0085 | 0.0014105 | 20.6319 | -0.0025724 | 0.99426 | 0.0018791 | -0.0050179 | 0.00090619 | 0.00071886 | -2.3402 | -0.12086 | 1.4803 | 2041.3131 | 1.0197
32 | Rinse |  | 5/29/2020 12:36:45 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.049 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.007457 | 8.6059 | 5.4038 | 0.08031 | 0.014134 | 0.022311 | 1.0419 | 0.15355 | 0.014143 | 1.3889 | 1.3113 | 0.0055704 | 0.013108 | -0.013076 | 0.16611 | 0.97831 | -0.018022 | 0.0044273 | 0.024107 | 0.23559 | 0.013906 | 0.99213 | 0.01189 | 0.020397 | 1.0244 | 0.0042948 | 0.054777 | 0.025252 | 1.0159 | 0.0043557 | 0.009237 | 0.017998 | 0.0054698 | 1.6643 | 2.5294 | 0.076399 | 13.7238 | 1.0419
33 | 2 ppb verification check |  | 5/29/2020 12:40:23 PM | Passed | C:\NexIONData\DataSet\052920\2 ppb verification check.050 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 2.2389 | -0.43588 | 2.2541 | 0.0071198 | 2.0282 | 2.0105 | 1.0203 | 2.1425 | 2.1395 | 202.7299 | 191.3976 | 2.0937 | 2.1016 | 2.1077 | 1.7147 | 0.97245 | 2.1937 | 2.1708 | 2.0846 | 221.6459 | 2.0114 | 0.97289 | 1.9023 | 2.0917 | 1.0064 | 2.0801 | 0.00079256 | 2.0814 | 0.98543 | 2.0962 | 2.1141 | 1.6022 | 2.1489 | 202.2203 | 208.4265 | 202.8817 | 282.9178 | 1.0203
34 | Rinse |  | 5/29/2020 12:44:04 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.051 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.027021 | 8.7792 | 6.6222 | 0.097576 | 0.029877 | 0.057252 | 1.0206 | 0.20471 | 0.04218 | 3.9116 | 3.6929 | 0.027076 | 0.042645 | 0.040673 | 0.46591 | 0.96264 | 0.033225 | 0.02734 | 0.055482 | 0.82437 | 0.044817 | 0.98355 | 0.11521 | 0.039798 | 1.0161 | 0.022995 | 0.069884 | 0.057125 | 1.0128 | 0.033942 | 0.059773 | 0.15224 | 0.027839 | 11.8186 | 5.6445 | 4.0276 | 18.8611 | 1.0206
35 | Sample blank |  | 5/29/2020 12:47:44 PM | Passed | C:\NexIONData\DataSet\052920\Sample blank.052 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0024318 | -0.75071 | -0.035974 | 0.0018298 | 0.00031077 | -0.010983 | 1.0203 | 0.014561 | -0.0056516 | -1.0047 | -0.94855 | 0.00040031 | 0.013023 | 0.052144 | -0.48184 | 0.97917 | 0.026215 | -0.0022449 | -0.0059967 | 0.016713 | 0.0021916 | 0.97422 | 0.018315 | 0.011875 | 1.0071 | -3.107e-05 | -0.0032776 | -0.0018151 | 0.99437 | 0.0011429 | -0.0054567 | 0.0020897 | -0.00011626 | -2.3325 | -0.038933 | 0.44123 | -1.1161 | 1.0203
36 | Rinse |  | 5/29/2020 12:51:25 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.053 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0060255 | 7.8422 | 4.8053 | 0.062316 | 0.01206 | 0.025001 | 1.0535 | 0.15053 | 0.013137 | 0.76034 | 0.71783 | 0.0058093 | 0.011614 | -0.0079795 | -0.16865 | 0.99105 | -0.033105 | 0.0046332 | -0.0068365 | 0.24756 | 0.010539 | 1.0079 | 0.012972 | 0.017413 | 1.0442 | 0.0037897 | 0.028106 | 0.019094 | 1.0448 | 0.004773 | 0.0084146 | 0.032582 | 0.0060365 | 1.6849 | 2.3528 | -0.81992 | 9.7795 | 1.0535
37 | Spike Blank |  | 5/29/2020 12:55:03 PM | Passed | C:\NexIONData\DataSet\052920\Spike Blank.054 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.9983 | 19.8356 | 23.2429 | 20.1556 | 20.8312 | 20.8004 | 1.0703 | 20.3543 | 20.9857 | 2100.9944 | 1983.5516 | 21.011 | 19.982 | 20.1821 | 20.7961 | 1.0152 | 21.2172 | 21.1563 | 20.6216 | 20.3563 | 19.6048 | 1.0023 | 19.8577 | 19.9855 | 1.0506 | 19.9616 | 19.6221 | 19.8302 | 1.0143 | 20.5063 | 20.5719 | 20.5922 | 21.153 | 2067.6056 | 2071.6541 | 2019.2891 | 2033.2241 | 1.0703
38 | Rinse |  | 5/29/2020 12:58:43 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.055 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.069936 | 8.1065 | 7.4216 | 0.15039 | 0.06459 | 0.079353 | 1.0573 | 0.16541 | 0.098085 | 8.047 | 7.5971 | 0.074252 | 0.10577 | 0.12906 | 1.5166 | 0.9947 | 0.075653 | 0.078809 | 0.084906 | 1.8167 | 0.12102 | 1.0086 | 0.24199 | 0.084189 | 1.05 | 0.066338 | 0.16523 | 0.12496 | 1.0497 | 0.081803 | 0.15701 | 0.31009 | 0.078544 | 25.8421 | 11.5379 | 8.924 | 29.6849 | 1.0573
39 | Rinse |  | 5/29/2020 1:02:22 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.056 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0079963 | 7.6462 | 4.7438 | 0.069715 | 4.6526e-05 | 0.012585 | 1.0498 | 0.10928 | 0.01437 | 0.9795 | 0.92474 | 0.0067061 | 0.013335 | -0.010178 | -0.14894 | 0.98382 | 0.031653 | 0.0049851 | 0.039991 | 0.22933 | 0.011732 | 1.0111 | 0.036664 | 0.018662 | 1.0546 | 0.0060012 | 0.025869 | 0.018807 | 1.0406 | 0.0062753 | 0.0075991 | 0.054023 | 0.0073409 | 1.5805 | 2.4719 | 0.13515 | 9.0263 | 1.0498
40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99.057 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.1809 | -14.4855 | 100.632 | 4.1753 | -0.45446 | -0.21126 | 1.0142 | 2.2854 | 7.6922 | -84.1389 | -79.4357 | -0.017029 | -0.51622 | 10.2296 | -7.516 | 0.97897 | -1.5515 | -0.035806 | 1.2833 | 162.9652 | 1.6463 | 0.95993 | -0.23737 | 0.71866 | 1.0135 | -0.050041 | -0.68597 | 69.4473 | 0.98583 | 0.077014 | -0.29096 | -0.11939 | 3.6008 | 9973.3778 | 40740.6268 | 4124.3476 | 64254.888 | 1.0142
41 | Mi1+HL 200520A 1_99 Duplicate |  | 5/29/2020 1:20:27 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Duplicate.058 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.15333 | 15.5856 | 198.706 | 4.4144 | 0.015346 | 0.39704 | 1.0194 | 3.5155 | 10.0684 | 14.7417 | 13.9177 | 0.03706 | 0.60183 | 12.5849 | 41.7858 | 0.9915 | 0.15342 | -0.096333 | 0.3066 | 165.8742 | 1.8624 | 0.97647 | 2.7221 | 0.91401 | 1.0358 | 0.068314 | -0.51492 | 69.3479 | 1.001 | 0.082317 | 0.049944 | 0.75813 | 3.7958 | 10056.6117 | 41492.6363 | 4161.3772 | 64816.485 | 1.0194
42 | Mi1+HL 200520A 1_9 |  | 5/29/2020 1:24:06 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_9.059 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.021764 | 33.2035 | 22.0234 | 1.43 | 0.15553 | 0.22699 | 1.0973 | 1.5756 | 8.3937 | 46.9827 | 44.3565 | -0.0055162 | 0.69033 | 12.1623 | 36.7679 | 1.0122 | 0.54361 | 0.036635 | 0.62381 | 160.5476 | 1.7343 | 0.99448 | 0.49559 | 0.093538 | 1.0701 | 0.024667 | 0.032788 | 62.7932 | 1.0412 | 0.032975 | 0.44785 | 0.20842 | 3.6378 | 10043.906 | 40000.5964 | 4020.6851 | 62239.278 | 1.0973
43 | Mi1+HL 200520A |  | 5/29/2020 1:27:46 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.060 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0024714 | 31.5492 | 3.4521 | 0.99786 | 0.10353 | 0.20005 | 1.1 | 0.93224 | 8.6508 | 19.1725 | 18.1008 | -0.038949 | 0.09079 | 12.0114 | 39.5073 | 0.92093 | 0.15112 | 0.053586 | 0.23547 | 188.1734 | 1.9039 | 0.88194 | 0.048257 | 0.014466 | 0.96601 | 0.017029 | 0.058216 | 64.6174 | 0.96198 | 0.014034 | 0.46066 | 0.031411 | 3.8266 | S | S | 4012.1251 | 62391.0665 | 1.1
44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.061 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.00075765 | 30.5901 | 2.8611 | 0.99045 | 0.12841 | 0.18842 | 1.1182 | 0.93133 | 8.5006 | 20.31 | 19.1747 | -0.036111 | 0.14547 | 11.9127 | 38.8046 | 0.92802 | 0.098502 | 0.056937 | 0.30557 | 186.6863 | 1.8445 | 0.90108 | 0.028913 | 0.010731 | 0.99829 | 0.015438 | 0.052147 | 61.886 | 0.97145 | 0.013019 | 0.44856 | 0.018073 | 3.6999 | 9672.4492 | S | 3909.0353 | 60836.6124 | 1.1182
45 | Rinse |  | 5/29/2020 1:35:06 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.062 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0026397 | 6.9305 | 4.0785 | 0.055445 | 0.019447 | 0.052996 | 1.0411 | 0.29508 | 0.015642 | 5.4017 | 5.0998 | 0.0031164 | 0.0096037 | 0.0034768 | -0.26135 | 0.96771 | 0.019633 | 0.0078027 | 0.043284 | 0.21797 | 0.0068062 | 0.99915 | 0.0070323 | 0.01541 | 1.0224 | 0.0013923 | 0.014933 | 0.042504 | 1.0269 | 0.0018666 | 0.0011744 | 0.029751 | 0.0051989 | 5.1101 | 22.4197 | 2.0488 | 40.0434 | 1.0411
46 | Mi1+HL 200520A 1_99 Spike |  | 5/29/2020 1:38:45 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Spike.063 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.0315 | 18.4155 | 22.337 | 19.8196 | 20.6693 | 20.5117 | 1.0375 | 19.6241 | 20.579 | 2072.224 | 1956.3894 | 20.7167 | 19.9348 | 20.5114 | 20.9299 | 0.98534 | 21.1202 | 21.0252 | 20.5165 | 21.272 | 19.4396 | 0.96334 | 19.8403 | 19.9352 | 1.0196 | 19.8953 | 19.5138 | 21.5703 | 1.0008 | 20.1629 | 20.3751 | 20.1956 | 20.8434 | 2110.6696 | 2400.5851 | 2043.9379 | 2579.7329 | 1.0375
47 | Rinse |  | 5/29/2020 1:42:26 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.064 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.01905 | 6.8231 | 5.4822 | 0.081263 | 0.017977 | 0.044549 | 1.024 | 0.19496 | 0.029867 | 3.7127 | 3.5052 | 0.017671 | 0.025633 | 0.021544 | -0.034988 | 0.96383 | 0.011468 | 0.022209 | 0.038445 | 0.35304 | 0.031506 | 0.99824 | 0.12793 | 0.027561 | 1.0283 | 0.016299 | 0.051547 | 0.050375 | 1.0351 | 0.020176 | 0.027256 | 0.15286 | 0.020104 | 6.1973 | 10.7785 | 2.8549 | 21.4814 | 1.024

## Con rev
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (mg/L) | Mg 24 (ug/L) | In 115 (ug/L) | U 238 (mg/L) | CeO 156 (mg/L) | Ce 140 (IS) | Ce++ 70 (mg/L) | Bkgd 220 (ug/L)
1 | Daily Performance Check |  | 5/29/2020 10:21:28 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.018 | C:\NexIONData\Method\Daily Performance.mth
2 | Daily Performance Check |  | 5/29/2020 10:23:57 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.019 | C:\NexIONData\Method\Daily Performance.mth
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (ug/L) | B 10 (ug/L) | Al 27 (ug/L) | Ti 47 (ug/L) | V-1 51 (ug/L) | V 51 (ug/L) | Sc 45 (IS) | Cr 52 (ug/L) | Mn 55 (ug/L) | Fe-1 57 (mg/L) | Fe-2 57 (ug/L) | Co 59 (ug/L) | Ni 60 (ug/L) | Cu 65 (ug/L) | Zn 66 (ug/L) | Ge 72 (IS) | As 75 (ug/L) | As-1 75 (ug/L) | Se 82 (ug/L) | Sr 87 (ug/L) | Mo 98 (ug/L) | Rh 103 (IS) | Ag 107 (ug/L) | Cd 111 (ug/L) | In 115 (IS) | Sb 121 (ug/L) | Sn 118 (ug/L) | Ba 137 (ug/L) | Tb 159 (IS) | Tl 205 (ug/L) | Pb 208 (ug/L) | Th 232 (ug/L) | U 238 (ug/L) | Na 23 (ug/L) | Mg 24 (ug/L) | K 39 (ug/L) | Ca 44 (ug/L) | Sc-1 45 (IS) | Cl 35 (ug/L)
40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99.057 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.1809 | -14.4855 | 100.632 | 4.1753 | -0.45446 | -0.21126 | 1.0142 | 2.2854 | 7.6922 | -84.1389 | -79.4357 | -0.017029 | -0.51622 | 10.2296 | -7.516 | 0.97897 | -1.5515 | -0.035806 | 1.2833 | 162.9652 | 1.6463 | 0.95993 | -0.23737 | 0.71866 | 1.0135 | -0.050041 | -0.68597 | 69.4473 | 0.98583 | 0.077014 | -0.29096 | -0.11939 | 3.6008 | 9973.3778 | 40740.6268 | 4124.3476 | 64254.888 | 1.0142
41 | Mi1+HL 200520A 1_99 Duplicate |  | 5/29/2020 1:20:27 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Duplicate.058 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.15333 | 15.5856 | 198.706 | 4.4144 | 0.015346 | 0.39704 | 1.0194 | 3.5155 | 10.0684 | 14.7417 | 13.9177 | 0.03706 | 0.60183 | 12.5849 | 41.7858 | 0.9915 | 0.15342 | -0.096333 | 0.3066 | 165.8742 | 1.8624 | 0.97647 | 2.7221 | 0.91401 | 1.0358 | 0.068314 | -0.51492 | 69.3479 | 1.001 | 0.082317 | 0.049944 | 0.75813 | 3.7958 | 10056.6117 | 41492.6363 | 4161.3772 | 64816.485 | 1.0194
42 | Mi1+HL 200520A 1_9 |  | 5/29/2020 1:24:06 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_9.059 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.021764 | 33.2035 | 22.0234 | 1.43 | 0.15553 | 0.22699 | 1.0973 | 1.5756 | 8.3937 | 46.9827 | 44.3565 | -0.0055162 | 0.69033 | 12.1623 | 36.7679 | 1.0122 | 0.54361 | 0.036635 | 0.62381 | 160.5476 | 1.7343 | 0.99448 | 0.49559 | 0.093538 | 1.0701 | 0.024667 | 0.032788 | 62.7932 | 1.0412 | 0.032975 | 0.44785 | 0.20842 | 3.6378 | 10043.906 | 40000.5964 | 4020.6851 | 62239.278 | 1.0973
43 | Mi1+HL 200520A |  | 5/29/2020 1:27:46 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.060 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0024714 | 31.5492 | 3.4521 | 0.99786 | 0.10353 | 0.20005 | 1.1 | 0.93224 | 8.6508 | 19.1725 | 18.1008 | -0.038949 | 0.09079 | 12.0114 | 39.5073 | 0.92093 | 0.15112 | 0.053586 | 0.23547 | 188.1734 | 1.9039 | 0.88194 | 0.048257 | 0.014466 | 0.96601 | 0.017029 | 0.058216 | 64.6174 | 0.96198 | 0.014034 | 0.46066 | 0.031411 | 3.8266 | S | S | 4012.1251 | 62391.0665 | 1.1
44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.061 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.00075765 | 30.5901 | 2.8611 | 0.99045 | 0.12841 | 0.18842 | 1.1182 | 0.93133 | 8.5006 | 20.31 | 19.1747 | -0.036111 | 0.14547 | 11.9127 | 38.8046 | 0.92802 | 0.098502 | 0.056937 | 0.30557 | 186.6863 | 1.8445 | 0.90108 | 0.028913 | 0.010731 | 0.99829 | 0.015438 | 0.052147 | 61.886 | 0.97145 | 0.013019 | 0.44856 | 0.018073 | 3.6999 | 9672.4492 | S | 3909.0353 | 60836.6124 | 1.1182
44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.061 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.00075765 | 30.5901 | 2.8611 | 0.99045 | 0.12841 | 0.18842 | 1.1182 | 0.93133 | 8.5006 | 20.31 | 19.1747 | -0.036111 | 0.14547 | 11.9127 | 38.8046 | 0.92802 | 0.098502 | 0.056937 | 0.30557 | 160.5476 | 1.8445 | 0.90108 | 0.028913 | 0.010731 | 0.99829 | 0.015438 | 0.052147 | 61.886 | 0.97145 | 0.013019 | 0.44856 | 0.018073 | 3.6999 | 10043.906 | 40000.5964 | 3909.0353 | 62239.278 | 1.1182

## Results
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (mg/L) | Mg 24 (ug/L) | In 115 (ug/L) | U 238 (mg/L) | CeO 156 (mg/L) | Ce 140 (IS) | Ce++ 70 (mg/L) | Bkgd 220 (ug/L)
1 | Daily Performance Check |  | 5/29/2020 10:21:28 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.018 | C:\NexIONData\Method\Daily Performance.mth
2 | Daily Performance Check |  | 5/29/2020 10:23:57 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.019 | C:\NexIONData\Method\Daily Performance.mth
Sample Id | Acquisition Time | Method File | Be 9 (ug/L) | B 10 (ug/L) | Al 27 (ug/L) | Ti 47 (ug/L) | V-1 51 (ug/L) | Cr 52 (ug/L) | Mn 55 (ug/L) | Fe-1 57 (mg/L) | Co 59 (ug/L) | Ni 60 (ug/L) | Cu 65 (ug/L) | Zn 66 (ug/L) | As-1 75 (ug/L) | Se 82 (ug/L) | Sr 87 (ug/L) | Mo 98 (ug/L) | Ag 107 (ug/L) | Cd 111 (ug/L) | Sb 121 (ug/L) | Sn 118 (ug/L) | Ba 137 (ug/L) | Tb 159 (IS) | Tl 205 (ug/L) | Pb 208 (ug/L) | Th 232 (ug/L) | U 238 (ug/L) | Na 23 (ug/L) | Mg 24 (ug/L) | K 39 (ug/L) | Ca 44 (ug/L)
Mi1+HL 200520A | 5/29/2020 1:31:26 PM | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.00075765 | 30.5901 | 2.8611 | 0.99045 | 0.12841 | 0.93133 | 8.5006 | 20.31 | -0.036111 | 0.14547 | 11.9127 | 38.8046 | 0.056937 | 0.30557 | 160.5476 | 1.8445 | 0.028913 | 0.010731 | 0.015438 | 0.052147 | 61.886 |  | 0.013019 | 0.44856 | 0.018073 | 3.6999 | 10043.906 | 40000.5964 | 3909.0353 | 62239.278

## QC review
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (ug/L) | B 10 (ug/L) | Al 27 (ug/L) | Ti 47 (ug/L) | V-1 51 (ug/L) | V 51 (ug/L) | Sc 45 (IS) | Cr 52 (ug/L) | Mn 55 (ug/L) | Fe-1 57 (mg/L) | Fe-2 57 (ug/L) | Co 59 (ug/L) | Ni 60 (ug/L) | Cu 65 (ug/L) | Zn 66 (ug/L) | Ge 72 (IS) | As 75 (ug/L) | As-1 75 (ug/L) | Se 82 (ug/L) | Sr 87 (ug/L) | Mo 98 (ug/L) | Rh 103 (IS) | Ag 107 (ug/L) | Cd 111 (ug/L) | In 115 (IS) | Sb 121 (ug/L) | Sn 118 (ug/L) | Ba 137 (ug/L) | Tb 159 (IS) | Tl 205 (ug/L) | Pb 208 (ug/L) | Th 232 (ug/L) | U 238 (ug/L) | Na 23 (ug/L) | Mg 24 (ug/L) | K 39 (ug/L) | Ca 44 (ug/L) | Sc-1 45 (IS) | Cl 35 (ug/L)
24 | 20ppb verification std |  | 5/29/2020 12:06:31 PM | Passed | C:\NexIONData\DataSet\052920\20ppb verification std.041 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.5855 | -1.1908 | 20.6169 | 0.037318 | 20.9746 | 21.1279 | 1.0595 | 21.1545 | 21.0357 | 2031.1531 | 1917.6143 | 20.7356 | 19.9582 | 20.448 | 20.6405 | 1.0028 | 21.6631 | 21.3196 | 20.787 | 2170.7668 | 19.7069 | 0.99599 | 20.0917 | 20.2229 | 1.0295 | 20.3646 | -0.0032876 | 20.6201 | 1.0301 | 20.1902 | 20.1736 | 20.1351 | 20.5208 | 1986.5761 | 2018.6356 | 2003.7628 | 3236.5155 | 1.0595 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
average | [6 empty] | 20.5855 | -1.1908 | 20.6169 | 0.037318 | 20.9746 | 21.1279 | 1.0595 | 21.1545 | 21.0357 | 2031.1531 | 1917.6143 | 20.7356 | 19.9582 | 20.448 | 20.6405 | 1.0028 | 21.6631 | 21.3196 | 20.787 | 2170.7668 | 19.7069 | 0.99599 | 20.0917 | 20.2229 | 1.0295 | 20.3646 | -0.0032876 | 20.6201 | 1.0301 | 20.1902 | 20.1736 | 20.1351 | 20.5208 | 1986.5761 | 2018.6356 | 2003.7628 | 3236.5155 | 1.0595 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
% recovery 90-110% | [5 empty] | Pass | 102.9% | -5.954% | 103.1% | 0.1866% | 104.9% | 105.6% | 5.297% | 105.8% | 105.2% | 101.6% | 95.88% | 103.7% | 99.79% | 102.2% | 103.2% | 5.014% | 108.3% | 106.6% | 103.9% | 108.5% | 98.53% | 4.98% | 100.5% | 101.1% | 5.147% | 101.8% | -0.01644% | 103.1% | 5.151% | 101% | 100.9% | 100.7% | 102.6% | 99.33% | 100.9% | 100.2% | 161.8%
25 | Ca+B+Sn+Ti check |  | 5/29/2020 12:10:25 PM | Passed | C:\NexIONData\DataSet\052920\Ca+B+Sn+Ti check.042 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.016314 | 19.7907 | 0.054748 | 21.3486 | -0.0086526 | 0.0085047 | 1.01 | 0.029981 | 0.0096204 | 1.2049 | 1.1375 | 0.012408 | 0.11415 | 0.022066 | -0.46537 | 0.97871 | 0.046883 | 0.00048143 | 0.053774 | 1.2355 | 0.017939 | 0.97135 | 0.099097 | 0.041625 | 1.0024 | 0.014683 | 21.2194 | 0.013705 | 0.98878 | 0.016935 | 0.0069514 | 0.039531 | 0.011865 | 0.85384 | 1.25 | 4.7883 | 2098.2652 | 1.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
average | [7 empty] | 19.7907 |  | 21.3486 | [22 empty] | 21.2194 | [9 empty] | 2098.2652
% recovery 85-115% | [5 empty] | Pass |  | 98.95% |  | 106.7% | [22 empty] | 106.1% | [9 empty] | 104.9%
29 | CRM |  | 5/29/2020 12:25:47 PM | Failed | C:\NexIONData\DataSet\052920\CRM.046 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.1146 | 0.81321 | 127.7002 | 0.087407 | 31.1346 | 31.3357 | 1.0913 | 21.5262 | 43.8148 | 94.7094 | 89.4153 | 26.9746 | 58.9908 | 19.8002 | 69.0159 | 0.96368 | 77.6086 | 81.8135 | 9.9844 | 291.1301 | 113.2625 | 0.95593 | 1.8741 | 9.6859 | 1.0258 | 9.7034 | 0.016441 | 48.8205 | 1.0106 | 9.6549 | 38.519 | 0.19583 | 10.3229 | 6038.2467 | 9062.1223 | 2483.0973 | 34263.7294 | 1.0913 | 0 | 0 | 0 | 0
average |  | #DIV/0! |  |  |  |  | 20.1146 | 0.81321 | 127.7002 | 0.087407 | 31.1346 | 31.3357 | 1.0913 | 21.5262 | 43.8148 | 94.7094 | 89.4153 | 26.9746 | 58.9908 | 19.8002 | 69.0159 | 0.96368 | 77.6086 | 81.8135 | 9.9844 | 291.1301 | 113.2625 | 0.95593 | 1.8741 | 9.6859 | 1.0258 | 9.7034 | 0.016441 | 48.8205 | 1.0106 | 9.6549 | 38.519 | 0.19583 | 10.3229 | 6038.2467 | 9062.1223 | 2483.0973 | 34263.7294 | 1.0913
target |  | 20 |  |  |  |  | 20 |  | 120 |  | 30 |  |  | 20 | 40 | 100 | 100 | 25 | 60 | 20 | 70 |  |  | 80 | 10 | 250 | 100 |  | 2 | 10 |  |  |  | 50 |  | 10 | 40 |  | 10 | 6000 | 9000 | 2500 | 35000
% recovery 85-115 |  | #DIV/0! |  |  |  | Pass | 100.6% | #DIV/0! | 106.4% | #DIV/0! | 103.8% | #DIV/0! | #DIV/0! | 107.6% | 109.5% | 94.71% | 89.42% | 107.9% | 98.32% | 99% | 98.59% | #DIV/0! | #DIV/0! | 102.3% | 99.84% | 116.5% | 113.3% | #DIV/0! | 93.7% | 96.86% | #DIV/0! | #DIV/0! | #DIV/0! | 97.64% | #DIV/0! | 96.55% | 96.3% | #DIV/0! | 103.2% | 100.6% | 100.7% | 99.32% | 97.9% | 0.003118% | #DIV/0! | #DIV/0!
28 | CRM 1_3 |  | 5/29/2020 12:22:07 PM | Passed | C:\NexIONData\DataSet\052920\CRM 1_3.045 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 5.6746 | -0.82135 | 32.3256 | 0.02744 | 7.4678 | 7.5102 | 1.0233 | 5.2232 | 10.1262 | 23.4619 | 22.1504 | 6.3331 | 15.1205 | 5.1957 | 20.3411 | 0.97779 | 21.6361 | 22.6704 | 3.0283 | 63.6926 | 24.6748 | 0.95987 | 0.44313 | 2.7524 | 1.0018 | 2.6125 | 0.00088994 | 12.5728 | 0.96652 | 2.5821 | 10.4265 | 0.1203 | 2.6568 | 1530.5736 | 2307.6853 | 627.9841 | 8687.4921 | 1.0233 | 0 | 0 | 0
average | [6 empty] | 5.6746 | -0.82135 | 32.3256 | 0.02744 | 7.4678 | 7.5102 | 1.0233 | 5.2232 | 10.1262 | 23.4619 | 22.1504 | 6.3331 | 15.1205 | 5.1957 | 20.3411 | 0.97779 | 21.6361 | 22.6704 | 3.0283 | 63.6926 | 24.6748 | 0.95987 | 0.44313 | 2.7524 | 1.0018 | 2.6125 | 0.00088994 | 12.5728 | 0.96652 | 2.5821 | 10.4265 | 0.1203 | 2.6568 | 1530.5736 | 2307.6853 | 627.9841 | 8687.4921 | 1.0233
target | [6 empty] | 5 |  | 30 |  | 7.5 |  |  | 5 | 10 | 25 | 25 | 6.25 | 15 | 5 | 17.5 |  |  | 20 | 2.5 | 62.5 | 25 |  | 0.5 | 2.5 |  |  |  | 12.5 |  | 2.5 | 10 |  | 2.5 | 1500 | 2250 | 625 | 8750
% recovery 85-115 | [5 empty] | Pass | 113.5% | #DIV/0! | 107.8% | #DIV/0! | 99.57% | #DIV/0! | #DIV/0! | 104.5% | 101.3% | 93.85% | 88.6% | 101.3% | 100.8% | 103.9% | 116.2% | #DIV/0! | #DIV/0! | 113.4% | 121.1% | 101.9% | 98.7% | #DIV/0! | 88.63% | 110.1% | #DIV/0! | #DIV/0! | #DIV/0! | 100.6% | #DIV/0! | 103.3% | 104.3% | #DIV/0! | 106.3% | 102% | 102.6% | 100.5% | 99.29% | #DIV/0!
35 | Sample blank |  | 5/29/2020 12:47:44 PM | Passed | C:\NexIONData\DataSet\052920\Sample blank.052 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0024318 | -0.75071 | -0.035974 | 0.0018298 | 0.00031077 | -0.010983 | 1.0203 | 0.014561 | -0.0056516 | -1.0047 | -0.94855 | 0.00040031 | 0.013023 | 0.052144 | -0.48184 | 0.97917 | 0.026215 | -0.0022449 | -0.0059967 | 0.016713 | 0.0021916 | 0.97422 | 0.018315 | 0.011875 | 1.0071 | -3.107e-05 | -0.0032776 | -0.0018151 | 0.99437 | 0.0011429 | -0.0054567 | 0.0020897 | -0.00011626 | -2.3325 | -0.038933 | 0.44123 | -1.1161 | 1.0203 | 0 | 0 | 0
37 | Spike Blank |  | 5/29/2020 12:55:03 PM | Passed | C:\NexIONData\DataSet\052920\Spike Blank.054 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.9983 | 19.8356 | 23.2429 | 20.1556 | 20.8312 | 20.8004 | 1.0703 | 20.3543 | 20.9857 | 2100.9944 | 1983.5516 | 21.011 | 19.982 | 20.1821 | 20.7961 | 1.0152 | 21.2172 | 21.1563 | 20.6216 | 20.3563 | 19.6048 | 1.0023 | 19.8577 | 19.9855 | 1.0506 | 19.9616 | 19.6221 | 19.8302 | 1.0143 | 20.5063 | 20.5719 | 20.5922 | 21.153 | 2067.6056 | 2071.6541 | 2019.2891 | 2033.2241 | 1.0703 | 0 | 0 | 0
% recovery 85-115% | [6 empty] | 105% | 102.9% | 116.4% | 100.8% | 104.2% | 104.1% | 0.25% | 101.7% | 105% | 105.1% | 99.23% | 105.1% | 99.85% | 100.6% | 106.4% | 0.18% | 106% | 105.8% | 103.1% | 101.7% | 98.01% | 0.1404% | 99.2% | 99.87% | 0.2175% | 99.81% | 98.13% | 99.16% | 0.09957% | 102.5% | 102.9% | 103% | 105.8% | 103.5% | 103.6% | 100.9% | 101.7% | 0.25%
not required for customers
40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99.057 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.001809 | -0.14486 | 1.0063 | 0.041753 | -0.0045446 | -0.0021126 | 0 | 0.022854 | 0.076922 | -0.84139 | -0.79436 | -0.00017029 | -0.0051622 | 0.1023 | -0.07516 | 0 | -0.015515 | -0.00035806 | 0.012833 | 1.6297 | 0.016463 | 0 | -0.0023737 | 0.0071866 | 0 | -0.00050041 | -0.0068597 | 0.69447 | 0 | 0.00077014 | -0.0029096 | -0.0011939 | 0.036008 | 99.7338 | 407.4063 | 41.2435 | 642.5489 | 0 | 0 | 0 | 0
41 | Mi1+HL 200520A 1_99 Duplicate |  | 5/29/2020 1:20:27 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Duplicate.058 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0015333 | 0.15586 | 1.9871 | 0.044144 | 0.00015346 | 0.0039704 | 0 | 0.035155 | 0.10068 | 0.14742 | 0.13918 | 0.0003706 | 0.0060183 | 0.12585 | 0.41786 | 0 | 0.0015342 | -0.00096333 | 0.003066 | 1.6587 | 0.018624 | 0 | 0.027221 | 0.0091401 | 0 | 0.00068314 | -0.0051492 | 0.69348 | 0 | 0.00082317 | 0.00049944 | 0.0075813 | 0.037958 | 100.5661 | 414.9264 | 41.6138 | 648.1649 | 0 | 0 | 0 | 0
avg |  | #DIV/0! |  |  |  |  | 0.0016712 | 0.0055005 | 1.4967 | 0.042948 | -0.0021956 | 0.00092892 | 0 | 0.029005 | 0.088803 | -0.34699 | -0.32759 | 0.00010016 | 0.00042805 | 0.11407 | 0.17135 | 0 | -0.0069906 | -0.0006607 | 0.0079495 | 1.6442 | 0.017543 | 0 | 0.012424 | 0.0081633 | 0 | 9.1364e-05 | -0.0060045 | 0.69398 | 0 | 0.00079666 | -0.0012051 | 0.0031937 | 0.036983 | 100.1499 | 411.1663 | 41.4286 | 645.3569 | 0
std |  | #DIV/0! |  |  |  |  | 0.00019493 | 0.21263 | 0.69349 | 0.0016907 | 0.0033221 | 0.0043014 | 0 | 0.0086981 | 0.016803 | 0.69919 | 0.66011 | 0.00038246 | 0.0079058 | 0.016655 | 0.34862 | 0 | 0.012056 | 0.00042799 | 0.0069062 | 0.020569 | 0.0015279 | 0 | 0.020926 | 0.0013814 | 0 | 0.00083689 | 0.0012095 | 0.0007032 | 0 | 3.7501e-05 | 0.0024106 | 0.006205 | 0.0013788 | 0.58855 | 5.3175 | 0.26184 | 3.9711 | 0
%RSD (±25%) |  | #DIV/0! |  |  |  |  | 11.6642 | 3865.7696 | 46.3348 | 3.9365 | -151.3059 | 463.0516 | #DIV/0! | 29.9886 | 18.9214 | -201.5044 | -201.5044 | 381.868 | 1846.923 | 14.6004 | 203.4543 | #DIV/0! | -172.4578 | -64.7789 | 86.8767 | 1.251 | 8.7096 | #DIV/0! | 168.4418 | 16.9216 | #DIV/0! | 916.0049 | -20.1439 | 0.10133 | #DIV/0! | 4.7074 | -200.033 | 194.2899 | 3.7282 | 0.58767 | 1.2933 | 0.63203 | 0.61533 | #DIV/0!
%RPD (±25%) |  | #VALUE! |  |  |  | Pass | 16.4957 | 5467.0238 | 65.5272 | 5.5671 | -213.9788 | 654.8538 | #DIV/0! | 42.4103 | 26.759 | -284.9703 | -284.9703 | 540.0429 | 2611.9435 | 20.6481 | 287.7278 | #DIV/0! | -243.8922 | -91.6111 | 122.8622 | 1.7692 | 12.3173 | #DIV/0! | 238.2127 | 23.9307 | #DIV/0! | 1295.4266 | -28.4878 | 0.1433 | #DIV/0! | 6.6572 | -282.8894 | 274.7674 | 5.2725 | 0.83109 | 1.829 | 0.89382 | 0.87021 | #DIV/0!
40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99.057 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.001809 | -0.14486 | 1.0063 | 0.041753 | -0.0045446 | -0.0021126 | 0 | 0.022854 | 0.076922 | -0.84139 | -0.79436 | -0.00017029 | -0.0051622 | 0.1023 | -0.07516 | 0 | -0.015515 | -0.00035806 | 0.012833 | 1.6297 | 0.016463 | 0 | -0.0023737 | 0.0071866 | 0 | -0.00050041 | -0.0068597 | 0.69447 | 0 | 0.00077014 | -0.0029096 | -0.0011939 | 0.036008 | 99.7338 | 407.4063 | 41.2435 | 642.5489 | 0 | 0 | 0
46 | Mi1+HL 200520A 1_99 Spike |  | 5/29/2020 1:38:45 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Spike.063 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.0315 | 18.4155 | 22.337 | 19.8196 | 20.6693 | 20.5117 | 1.0375 | 19.6241 | 20.579 | 2072.224 | 1956.3894 | 20.7167 | 19.9348 | 20.5114 | 20.9299 | 0.98534 | 21.1202 | 21.0252 | 20.5165 | 21.272 | 19.4396 | 0.96334 | 19.8403 | 19.9352 | 1.0196 | 19.8953 | 19.5138 | 21.5703 | 1.0008 | 20.1629 | 20.3751 | 20.1956 | 20.8434 | 2110.6696 | 2400.5851 | 2043.9379 | 2579.7329 | 1.0375 | 0 | 0
% recovery 70-130% |  | #VALUE! |  |  |  | Pass | 100.1% | 92.8% | 106.7% | 98.89% | 103.4% | 102.6% | 5.188% | 98.01% | 102.5% | 103.7% | 97.86% | 103.6% | 99.7% | 102% | 105% | 4.927% | 105.7% | 105.1% | 102.5% | 98.21% | 97.12% | 4.817% | 99.21% | 99.64% | 5.098% | 99.48% | 97.6% | 104.4% | 5.004% | 100.8% | 101.9% | 101% | 104% | 100.5% | 99.66% | 100.1% | 96.86% | 0.05188%
43 | Mi1+HL 200520A |  | 5/29/2020 1:27:46 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.060 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0024714 | 31.5492 | 3.4521 | 0.99786 | 0.10353 | 0.20005 | 1.1 | 0.93224 | 8.6508 | 19.1725 | 18.1008 | -0.038949 | 0.09079 | 12.0114 | 39.5073 | 0.92093 | 0.15112 | 0.053586 | 0.23547 | 188.1734 | 1.9039 | 0.88194 | 0.048257 | 0.014466 | 0.96601 | 0.017029 | 0.058216 | 64.6174 | 0.96198 | 0.014034 | 0.46066 | 0.031411 | 3.8266 | S | S | 4012.1251 | 62391.0665 | 1.1 | 0 | 0 | 0
44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.061 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.00075765 | 30.5901 | 2.8611 | 0.99045 | 0.12841 | 0.18842 | 1.1182 | 0.93133 | 8.5006 | 20.31 | 19.1747 | -0.036111 | 0.14547 | 11.9127 | 38.8046 | 0.92802 | 0.098502 | 0.056937 | 0.30557 | 186.6863 | 1.8445 | 0.90108 | 0.028913 | 0.010731 | 0.99829 | 0.015438 | 0.052147 | 61.886 | 0.97145 | 0.013019 | 0.44856 | 0.018073 | 3.6999 | 9672.4492 | S | 3909.0353 | 60836.6124 | 1.1182 | 0 | 0 | 0
avg |  | #DIV/0! |  |  |  |  | 0.0016145 | 31.0697 | 3.1566 | 0.99416 | 0.11597 | 0.19424 | 1.1091 | 0.93178 | 8.5757 | 19.7413 | 18.6377 | -0.03753 | 0.11813 | 11.962 | 39.156 | 0.92448 | 0.12481 | 0.055262 | 0.27052 | 187.4298 | 1.8742 | 0.89151 | 0.038585 | 0.012599 | 0.98215 | 0.016233 | 0.055182 | 63.2517 | 0.96672 | 0.013527 | 0.45461 | 0.024742 | 3.7632 | 9672.4492 | #DIV/0! | 3960.5802 | 61613.8395 | 1.1091
std |  | #DIV/0! |  |  |  |  | 0.0012118 | 0.67821 | 0.41791 | 0.0052385 | 0.017592 | 0.0082255 | 0.012874 | 0.0006399 | 0.10624 | 0.80435 | 0.75938 | 0.0020068 | 0.038665 | 0.069761 | 0.4969 | 0.0050151 | 0.037204 | 0.0023698 | 0.049569 | 1.0515 | 0.04195 | 0.013531 | 0.013678 | 0.0026411 | 0.022826 | 0.0011247 | 0.0042912 | 1.9314 | 0.0067029 | 0.00071787 | 0.0085574 | 0.0094312 | 0.089577 | #DIV/0! | #DIV/0! | 72.8955 | 1099.165 | 0.012874
%RSD (±25%) |  | #DIV/0! |  |  |  |  | 75.0561 | 2.1829 | 13.2393 | 0.52693 | 15.1693 | 4.2348 | 1.1608 | 0.068674 | 1.2389 | 4.0744 | 4.0744 | -5.3472 | 32.7306 | 0.58319 | 1.269 | 0.54248 | 29.8086 | 4.2884 | 18.3235 | 0.56102 | 2.2383 | 1.5178 | 35.4497 | 20.9634 | 2.3241 | 6.9282 | 7.7765 | 3.0536 | 0.69336 | 5.307 | 1.8824 | 38.1185 | 2.3803 | #DIV/0! | #DIV/0! | 1.8405 | 1.784 | 1.1608
%RPD (±25%) |  | #VALUE! |  |  |  | Pass | 106.1454 | 3.0871 | 18.7233 | 0.7452 | 21.4527 | 5.9889 | 1.6416 | 0.09712 | 1.752 | 5.7621 | 5.7621 | -7.562 | 46.2881 | 0.82475 | 1.7947 | 0.76718 | 42.1557 | 6.0647 | 25.9134 | 0.7934 | 3.1654 | 2.1465 | 50.1335 | 29.6468 | 3.2868 | 9.7979 | 10.9977 | 4.3184 | 0.98057 | 7.5053 | 2.6621 | 53.9076 | 3.3663 | #VALUE! | #VALUE! | 2.6029 | 2.5229 | 1.6416
117 | 2 ppb verification check |  | 1/9/2020 3:04:55 PM | Passed | C:\NexIONData\DataSet\010920\2 ppb verification check.059 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 2.243 | 1.9702 | 2.8029 | 0.0047002 | 2.0218 | 2.0427 | 0.9541 | 2.0596 | 2.0228 | 200.18 | 196.9293 | 2.0436 | 2.0228 | 2.0137 | 1.1734 | 0.98018 | 2.0503 | 2.0232 | 2.0803 | 191.9781 | 1.9067 | 0.97564 | 2.0172 | 2.0706 | 0.96316 | 2.0174 | 3.5986e-05 | 2.1629 | 0.96544 | 2.0186 | 2.0297 | 1.8861 | 2.0324 | 201.5166 | 202.398 | 208.2098 | 284.4807 | 0.9541 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
average | [6 empty] | 2.243 | 1.9702 | 2.8029 | 0.0047002 | 2.0218 | 2.0427 | 0.9541 | 2.0596 | 2.0228 | 200.18 | 196.9293 | 2.0436 | 2.0228 | 2.0137 | 1.1734 | 0.98018 | 2.0503 | 2.0232 | 2.0803 | 191.9781 | 1.9067 | 0.97564 | 2.0172 | 2.0706 | 0.96316 | 2.0174 | 3.5986e-05 | 2.1629 | 0.96544 | 2.0186 | 2.0297 | 1.8861 | 2.0324 | 201.5166 | 202.398 | 208.2098 | 284.4807 | 0.9541 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
% recovery 90-110% | [5 empty] | Pass | 1.1215 | 0.98511 | 1.4015 | 0.0023501 | 1.0109 | 1.0214 | 0.47705 | 1.0298 | 1.0114 | 100.09 | 98.4646 | 1.0218 | 1.0114 | 1.0068 | 0.58672 | 0.49009 | 1.0252 | 1.0116 | 1.0402 | 95.989 | 0.95336 | 0.48782 | 1.0086 | 1.0353 | 0.48158 | 1.0087 | 1.7993e-05 | 1.0815 | 0.48272 | 1.0093 | 1.0148 | 0.94304 | 1.0162 | 100.7583 | 101.199 | 104.1049 | 142.2404
33 | 2 ppb verification check |  | 5/29/2020 12:40:23 PM | Passed | C:\NexIONData\DataSet\052920\2 ppb verification check.050 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 2.2389 | -0.43588 | 2.2541 | 0.0071198 | 2.0282 | 2.0105 | 1.0203 | 2.1425 | 2.1395 | 202.7299 | 191.3976 | 2.0937 | 2.1016 | 2.1077 | 1.7147 | 0.97245 | 2.1937 | 2.1708 | 2.0846 | 221.6459 | 2.0114 | 0.97289 | 1.9023 | 2.0917 | 1.0064 | 2.0801 | 0.00079256 | 2.0814 | 0.98543 | 2.0962 | 2.1141 | 1.6022 | 2.1489 | 202.2203 | 208.4265 | 202.8817 | 282.9178 | 1.0203 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
average | [6 empty] | 2.2389 | -0.43588 | 2.2541 | 0.0071198 | 2.0282 | 2.0105 | 1.0203 | 2.1425 | 2.1395 | 202.7299 | 191.3976 | 2.0937 | 2.1016 | 2.1077 | 1.7147 | 0.97245 | 2.1937 | 2.1708 | 2.0846 | 221.6459 | 2.0114 | 0.97289 | 1.9023 | 2.0917 | 1.0064 | 2.0801 | 0.00079256 | 2.0814 | 0.98543 | 2.0962 | 2.1141 | 1.6022 | 2.1489 | 202.2203 | 208.4265 | 202.8817 | 282.9178 | 1.0203 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
% recovery 85-115% | [5 empty] | Pass | 111.9% | -21.79% | 112.7% | 0.356% | 101.4% | 100.5% | 51.01% | 107.1% | 107% | 101.4% | 95.7% | 104.7% | 105.1% | 105.4% | 85.73% | 48.62% | 109.7% | 108.5% | 104.2% | 11082% | 100.6% | 48.64% | 95.11% | 104.6% | 50.32% | 104% | 0.03963% | 104.1% | 49.27% | 104.8% | 105.7% | 80.11% | 107.4% | 101.1% | 104.2% | 101.4% | 141.5% | 51.01% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%

## Concentrations RSDs
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 | Mg 24 | In 115 | U 238 | CeO 156 | Ce 140 (IS) | Ce++ 70 | Bkgd 220
1 | Daily Performance Check |  | 5/29/2020 10:21:28 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.018 | C:\NexIONData\Method\Daily Performance.mth
2 | Daily Performance Check |  | 5/29/2020 10:23:57 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.019 | C:\NexIONData\Method\Daily Performance.mth
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 | B 10 | Al 27 | Ti 47 | V-1 51 | V 51 | Sc 45 (IS) | Cr 52 | Mn 55 | Fe-1 57 | Fe-2 57 | Co 59 | Ni 60 | Cu 65 | Zn 66 | Ge 72 (IS) | As 75 | As-1 75 | Se 82 | Sr 87 | Mo 98 | Rh 103 (IS) | Ag 107 | Cd 111 | In 115 (IS) | Sb 121 | Sn 118 | Ba 137 | Tb 159 (IS) | Tl 205 | Pb 208 | Th 232 | U 238 | Na 23 | Mg 24 | K 39 | Ca 44 | Sc-1 45 (IS) | Cl 35
3 | Rinse |  | 5/29/2020 10:31:04 AM |  | C:\NexIONData\DataSet\052920\Rinse.020 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
4 | Rinse |  | 5/29/2020 10:37:14 AM |  | C:\NexIONData\DataSet\052920\Rinse.021 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
5 | Rinse |  | 5/29/2020 10:41:20 AM |  | C:\NexIONData\DataSet\052920\Rinse.022 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
6 | Rinse |  | 5/29/2020 10:44:59 AM |  | C:\NexIONData\DataSet\052920\Rinse.023 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
7 | Standard Blank |  | 5/29/2020 10:48:37 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.024 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
8 | Standard Blank |  | 5/29/2020 10:52:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.025 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
9 | 20 ppm Ca |  | 5/29/2020 10:55:53 AM |  | C:\NexIONData\DataSet\052920\20 ppm Ca .026 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
10 | 20 ppm Cl |  | 5/29/2020 10:59:32 AM |  | C:\NexIONData\DataSet\052920\20 ppm Cl.027 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
11 | Rinse |  | 5/29/2020 11:03:12 AM |  | C:\NexIONData\DataSet\052920\Rinse.028 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
12 | Rinse |  | 5/29/2020 11:06:50 AM |  | C:\NexIONData\DataSet\052920\Rinse.029 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
13 | Standard Blank |  | 5/29/2020 11:26:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.030 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
14 | Standard Blank |  | 5/29/2020 11:29:53 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.031 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
15 | std blank |  | 5/29/2020 11:33:31 AM |  | C:\NexIONData\DataSet\052920\std blank.032 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | [6 empty] | 0.975% | [8 empty] | 2.007% | [5 empty] | 2.569% |  |  | 1.386% |  |  |  | 0.529% | [8 empty] | 0.975%
16 | 1ppb |  | 5/29/2020 11:37:10 AM |  | C:\NexIONData\DataSet\052920\1ppb.033 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.8013% | 96.14% | 5.698% | 0.5534% | 0.6999% | 0.381% | 0.2184% | 2.672% | 0.7166% | 48.64% |  | 0.1696% | 0.6266% | 0.8601% | 4.492% | 1.507% | 6.771% | 1.517% | 6.3% | 1.196% | 0.4839% | 1.86% | 5.985% | 1.831% | 1.463% | 1.72% | 2.14% | 2.32% | 0.2627% | 0.9119% | 0.8963% | 8.757% | 1.31% |  |  |  |  | 0.2184%
17 | 10ppb |  | 5/29/2020 11:40:49 AM |  | C:\NexIONData\DataSet\052920\10ppb.034 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 3.963% | 1.92% | 3.933% | 3.535% | 3.725% | 4.077% | 4.24% | 2.674% | 2.779% | 9.829% |  | 3.591% | 3.415% | 3.948% | 4.021% | 2.742% | 4.068% | 3.056% | 3.267% | 2.668% | 3.896% | 4.399% | 3.332% | 6.481% | 5.107% | 5.859% | 4.968% | 4.925% | 3.802% | 1.963% | 2.579% | 1.581% | 2.401% |  |  |  |  | 4.24%
18 | 50ppb |  | 5/29/2020 11:44:29 AM |  | C:\NexIONData\DataSet\052920\50ppb.035 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 1.446% | 1.784% | 0.3824% | 1.191% | 0.9964% | 0.9319% | 1.49% | 1.775% | 1.726% | 0.09957% |  | 1.849% | 1.333% | 0.8769% | 1.429% | 0.9125% | 2.44% | 1.552% | 2.588% | 1.413% | 1.401% | 1.653% | 0.9084% | 0.2427% | 0.7325% | 0.3521% | 0.8801% | 0.9519% | 0.6654% | 2.023% | 0.8559% | 0.8339% | 1.494% |  |  |  |  | 1.49%
19 | 100ppb |  | 5/29/2020 11:48:09 AM |  | C:\NexIONData\DataSet\052920\100ppb.036 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.1893% | 2.465% | 0.9661% | 0.8189% | 1.537% | 1.604% | 0.1595% | 2.761% | 3.806% | 3.574% |  | 3.28% | 2.376% | 2.801% | 1.831% | 1.426% | 1.753% | 2.235% | 2.125% | 3.376% | 2.304% | 0.898% | 1.24% | 0.7646% | 0.3336% | 1.428% | 1.25% | 1.505% | 2.162% | 4.071% | 3.089% | 3.005% | 3.502% |  |  |  |  | 0.1595%
20 | 100 PPB |  | 5/29/2020 11:51:50 AM |  | C:\NexIONData\DataSet\052920\100 PPB.037 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 25.78% | 1.636% | 41.02% | 20.12% | 19.91% | 33.5% | 1.73% | 84.02% | 24.23% | 1.739% | 1.739% | 19.2% | 19.54% | 9.964% | 16.94% | 3.353% | 40.43% | 20.83% | 21.01% | 16.41% | 15.99% | 4.194% | 19.96% | 17.38% | 2.857% | 16.32% | 17.26% | 28.72% | 2.736% | 14.76% | 19.56% | 27.97% | 20.88% | 0.2356% | 0.6305% | 1.294% | 0.9403% | 1.73%
21 | 1000 PPB |  | 5/29/2020 11:55:30 AM |  | C:\NexIONData\DataSet\052920\1000 PPB.038 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 14.2% | 0.7197% | 87.53% | 17.97% | 9.751% | 39.65% | 2.596% | 176.3% | 7.154% | 1.963% | 1.963% | 4.15% | 6.788% | 13.7% | 14.89% | 2.528% | 247.1% | 11.94% | 169.9% | 14.41% | 5.246% | 1.762% | 6.272% | 33.3% | 2.63% | 7.943% | 19.3% | 11.77% | 0.935% | 6.225% | 2348% | 17.65% | 10.12% | 1.122% | 1.773% | 0.615% | 0.2172% | 2.596%
22 | 5000 PPB |  | 5/29/2020 11:59:08 AM |  | C:\NexIONData\DataSet\052920\5000 PPB.039 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 51.29% | 3.704% | 49.18% | 5.127% | 4.973% | 12.33% | 1.331% | 1.001% | 2.062% | 1.605% | 1.605% | 0.6527% | 2.057% | 1.877% | 8.37% | 0.5559% | 235.7% | 10.8% | 48.68% | 5.31% | 5.436% | 0.4984% | 5.354% | 18.77% | 0.8188% | 15.33% | 18.22% | 0.7945% | 0.6048% | 7.831% | 449.8% | 21.79% | 32.74% | 0.9994% | 0.2009% | 0.8373% | 0.9041% | 1.331%
23 | 10000 PPB |  | 5/29/2020 12:02:47 PM |  | C:\NexIONData\DataSet\052920\10000 PPB.040 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 50.33% | 2.564% | 65.09% | 3.451% | 6.935% | 7.294% | 0.8884% | 10.12% | 2.075% | 0.7487% | 0.7487% | 1.118% | 2.506% | 3.011% | 16.25% | 0.8836% | 15.52% | 11.19% | 84.21% | 6.753% | 1.612% | 0.2111% | 14.23% | 23.24% | 0.698% | 6.285% | 14.67% | 8.027% | 0.4971% | 10.47% | 15.17% | 29.05% | 37.28% | S | 1.002% | S | 0.3131% | 0.8884%
24 | 20ppb verification std |  | 5/29/2020 12:06:31 PM | Passed | C:\NexIONData\DataSet\052920\20ppb verification std.041 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 1.992% | 8.032% | 4.631% | 28.74% | 0.9043% | 1.791% | 0.8022% | 4.532% | 1.203% | 0.8687% | 0.8687% | 0.6243% | 0.4491% | 0.3594% | 0.2453% | 0.8177% | 1.28% | 0.9152% | 1.435% | 0.4224% | 0.7303% | 0.802% | 1.78% | 1.278% | 0.8791% | 1.481% | 41.18% | 0.7871% | 0.9265% | 0.8898% | 0.9147% | 1.798% | 0.6059% | 1.36% | 1.4% | 0.8385% | 0.5557% | 0.8022%
25 | Ca+B+Sn+Ti check |  | 5/29/2020 12:10:25 PM | Passed | C:\NexIONData\DataSet\052920\Ca+B+Sn+Ti check.042 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 29.03% | 4.49% | 493.7% | 3.686% | 34.18% | 118.1% | 4.39% | 158% | 25.35% | 57.61% | 57.61% | 16.52% | 8.359% | 9.35% | 3.638% | 4.693% | 84.92% | 613.2% | 53.38% | 11.33% | 7.978% | 3.701% | 19.27% | 11.99% | 4.098% | 15.66% | 2.892% | 30.69% | 3.564% | 3.957% | 13.69% | 3.629% | 10.02% | 6.082% | 26.2% | 31.88% | 2.951% | 4.39%
26 | Verification Check |  | 5/29/2020 12:14:49 PM | Passed | C:\NexIONData\DataSet\052920\Verification Check.043 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 2.371% | 12.84% | 2.655% | 34.45% | 1.351% | 1.259% | 2.743% | 1.649% | 0.6319% | 1.89% | 1.89% | 1.418% | 1.262% | 1.343% | 0.38% | 1.849% | 0.6579% | 1.138% | 0.5962% | 1.22% | 0.9082% | 1.144% | 0.3352% | 0.2598% | 1.241% | 0.08843% | 30.09% | 0.7425% | 2.383% | 1.312% | 2.341% | 1.381% | 2.01% | 2.287% | 0.3989% | 1.1% | 0.7832% | 2.743%
27 | Rinse |  | 5/29/2020 12:18:29 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.044 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 25.57% | 10.28% | 30.9% | 26.53% | 1578% | 48.34% | 1.962% | 69.8% | 29.22% | 48.4% | 48.4% | 36.03% | 33.53% | 347.6% | 975.1% | 2.556% | 173.2% | 18.81% | 93.74% | 29.39% | 27.59% | 3.612% | 17.94% | 30.11% | 3.885% | 32.35% | 16.4% | 41.41% | 6.275% | 31.23% | 52.2% | 8.33% | 37.81% | 33.46% | 35.67% | 36.39% | 22.14% | 1.962%
28 | CRM 1_3 |  | 5/29/2020 12:22:07 PM | Passed | C:\NexIONData\DataSet\052920\CRM 1_3.045 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.7817% | 7.027% | 0.458% | 19.54% | 2.356% | 2.467% | 1.589% | 2.378% | 1.779% | 2.898% | 2.898% | 0.9768% | 1.483% | 0.9931% | 0.9678% | 1.685% | 2.083% | 1.525% | 2.632% | 2.203% | 2.278% | 2.579% | 5.689% | 1.497% | 2.143% | 2.242% | 240.6% | 2.473% | 0.7494% | 0.5927% | 0.8337% | 3.25% | 1.148% | 0.5437% | 0.6625% | 1.124% | 1.898% | 1.589%
29 | CRM |  | 5/29/2020 12:25:47 PM | Failed | C:\NexIONData\DataSet\052920\CRM.046 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 2.566% | 85.28% | 2.245% | 11.46% | 1.879% | 2.747% | 1.03% | 6.265% | 0.8802% | 2.283% | 2.283% | 0.4687% | 0.8986% | 0.9422% | 0.6217% | 1.207% | 0.474% | 0.4934% | 1.016% | 0.713% | 2.265% | 0.9227% | 5.422% | 1.334% | 0.4134% | 1.532% | 17.95% | 1.379% | 1.956% | 2.53% | 2.671% | 7.66% | 3.363% | 2.463% | 2.612% | 1.712% | 1.141% | 1.03%
30 | Rinse |  | 5/29/2020 12:29:27 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.047 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 50.7% | 11.07% | 39.81% | 35.85% | 65.08% | 45.17% | 2.304% | 25.89% | 46.23% | 55.5% | 55.5% | 50.75% | 49.4% | 940.2% | 2720% | 2.049% | 118.9% | 52.49% | 168.9% | 49.71% | 39.46% | 1.782% | 21.28% | 23.03% | 1.927% | 51.8% | 27.8% | 45.63% | 2.32% | 44.21% | 60.3% | 24.67% | 60.31% | 56.29% | 48.75% | 83.35% | 42.55% | 2.304%
31 | B, Ca, Ti, Sn check |  | 5/29/2020 12:33:05 PM | Passed | C:\NexIONData\DataSet\052920\B, Ca, Ti, Sn check.048 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 22.01% | 2.209% | 256.2% | 0.8097% | 12.14% | 74.89% | 0.5617% | 28.6% | 36.4% | 239.1% | 239.1% | 25.87% | 6.646% | 57.1% | 32.64% | 0.3637% | 90.28% | 33.08% | 770.5% | 13.03% | 12.18% | 0.6342% | 6.429% | 11.31% | 1.203% | 12.88% | 0.9192% | 13.89% | 1.789% | 16.08% | 13.02% | 18.87% | 41.5% | 3.169% | 103.3% | 15.19% | 0.9244% | 0.5617%
32 | Rinse |  | 5/29/2020 12:36:45 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.049 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 77.62% | 9.268% | 30.46% | 26.69% | 30.63% | 16.34% | 0.898% | 10.56% | 44.94% | 49.48% | 49.48% | 71.48% | 64.52% | 59.75% | 94.27% | 0.9616% | 190.2% | 122.4% | 131.4% | 57.77% | 53.37% | 0.4269% | 27.3% | 57.84% | 0.8623% | 85.01% | 18.3% | 58.15% | 1.452% | 67.43% | 119.8% | 26.81% | 88.45% | 78.97% | 46.29% | 843.9% | 33.62% | 0.898%
33 | 2 ppb verification check |  | 5/29/2020 12:40:23 PM | Passed | C:\NexIONData\DataSet\052920\2 ppb verification check.050 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 1.243% | 11.61% | 5.321% | 30.17% | 0.7688% | 0.9803% | 1.113% | 0.572% | 0.2247% | 0.8021% | 0.8021% | 0.9426% | 2.431% | 0.6444% | 3.085% | 0.5122% | 2.848% | 0.526% | 2.799% | 1.53% | 0.6066% | 0.8662% | 5.056% | 0.4964% | 0.763% | 0.3639% | 231.8% | 1.117% | 0.4294% | 1.69% | 1.476% | 10.11% | 1.533% | 1.563% | 0.5811% | 0.4155% | 0.4751% | 1.113%
34 | Rinse |  | 5/29/2020 12:44:04 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.051 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 36.37% | 12.7% | 36.67% | 33.01% | 40.11% | 28.24% | 2.546% | 20.95% | 30.58% | 32.46% | 32.46% | 37.87% | 34.93% | 50.15% | 68.11% | 2.472% | 101.7% | 41.56% | 81.95% | 38.76% | 33.61% | 1.034% | 22.01% | 19.45% | 3.053% | 29.92% | 29.16% | 40.87% | 2.302% | 27.56% | 38.3% | 16.53% | 41.86% | 37.18% | 39.88% | 44.73% | 37.87% | 2.546%
35 | Sample blank |  | 5/29/2020 12:47:44 PM | Passed | C:\NexIONData\DataSet\052920\Sample blank.052 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 15.72% | 13.68% | 265.8% | 257.9% | 234.1% | 30.61% | 0.3505% | 100.2% | 23.24% | 61.49% | 61.49% | 69.81% | 14.62% | 10.74% | 11.99% | 0.1599% | 232% | 34.95% | 157.7% | 46.24% | 16.84% | 0.3009% | 9.064% | 44.42% | 1.28% | 403.4% | 46.29% | 165.5% | 1.162% | 21% | 16.3% | 20.5% | 47.83% | 2.432% | 209% | 31.21% | 6.76% | 0.3505%
36 | Rinse |  | 5/29/2020 12:51:25 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.053 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 64.38% | 6.153% | 26.2% | 27.59% | 34.67% | 5.079% | 0.7108% | 5.71% | 33.91% | 16.65% | 16.65% | 61.5% | 23.89% | 82.58% | 70.68% | 0.5303% | 121.9% | 38.17% | 190.8% | 44.51% | 53.67% | 0.8185% | 28.01% | 35.47% | 0.5531% | 88.7% | 22.07% | 40.1% | 0.6643% | 52.63% | 95.49% | 9.298% | 72.89% | 62.67% | 37.91% | 8.654% | 33.16% | 0.7108%
37 | Spike Blank |  | 5/29/2020 12:55:03 PM | Passed | C:\NexIONData\DataSet\052920\Spike Blank.054 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 1.138% | 2.66% | 2.751% | 1.606% | 1.225% | 2.4% | 1.701% | 6.117% | 1.306% | 1.887% | 1.887% | 1.464% | 1.044% | 0.8903% | 1.603% | 1.419% | 2.367% | 1.074% | 2.081% | 1.674% | 2.284% | 1.442% | 1.658% | 1.339% | 1.524% | 0.832% | 1.497% | 0.4542% | 2.312% | 1.3% | 2.036% | 2.171% | 1.5% | 0.2494% | 0.4359% | 0.446% | 0.9217% | 1.701%
38 | Rinse |  | 5/29/2020 12:58:43 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.055 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 36.16% | 8.706% | 24.25% | 23.53% | 35.3% | 28.31% | 0.4772% | 6.249% | 25.34% | 24.32% | 24.32% | 28.28% | 25.48% | 26.11% | 30.39% | 1.132% | 44.69% | 19.49% | 48.66% | 24.43% | 25.26% | 0.9141% | 22.59% | 23.77% | 1.133% | 28.32% | 20.56% | 26.76% | 0.882% | 27.21% | 26.39% | 18.1% | 28.02% | 26.19% | 28.62% | 28.54% | 26.19% | 0.4772%
39 | Rinse |  | 5/29/2020 1:02:22 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.056 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 73.9% | 7.987% | 31.03% | 19.77% | 9014% | 59.8% | 0.07776% | 17.99% | 26.53% | 16.13% | 16.13% | 52.98% | 35.96% | 87.82% | 149.7% | 0.6161% | 83.03% | 64.19% | 22.04% | 44.95% | 57.89% | 1.099% | 16.7% | 3.917% | 0.6689% | 46.03% | 26.46% | 34.94% | 1.255% | 41.55% | 99.4% | 7.671% | 57.23% | 61.53% | 40.17% | 137.8% | 41.55% | 0.07776%
40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99.057 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 59.12% | 197.3% | 28.77% | 36.71% | 63.49% | 337% | 1.997% | 64.36% | 3.323% | 60.62% | 60.62% | 71.41% | 37.41% | 7.384% | 26.88% | 2.003% | 111.6% | 390.3% | 55.17% | 2.282% | 3.178% | 2.252% | 13.05% | 49.18% | 2.296% | 47.9% | 2.376% | 2.49% | 1.49% | 29.93% | 27% | 14.06% | 1.627% | 1.543% | 1.309% | 3.039% | 1.211% | 1.997%
41 | Mi1+HL 200520A 1_99 Duplicate |  | 5/29/2020 1:20:27 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Duplicate.058 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 28.87% | 96.73% | 7.992% | 28.15% | 1007% | 237.6% | 2.475% | 67.91% | 2.182% | 258.4% | 258.4% | 106.6% | 26.6% | 5.531% | 3.434% | 2.17% | 2057% | 113.2% | 1384% | 2.277% | 7.038% | 1.781% | 31.16% | 4.701% | 1.17% | 8.855% | 11.54% | 1.039% | 0.4561% | 43.38% | 265.3% | 15.75% | 1.478% | 1.321% | 1.285% | 2.181% | 1.527% | 2.475%
42 | Mi1+HL 200520A 1_9 |  | 5/29/2020 1:24:06 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_9.059 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 54.26% | 13.37% | 25.56% | 9.594% | 11.19% | 14.16% | 2.099% | 13.31% | 2.054% | 10.59% | 10.59% | 77.38% | 5.285% | 2.245% | 1.373% | 0.8347% | 98.06% | 23.03% | 72.48% | 2.08% | 3.218% | 1.41% | 27.24% | 46.08% | 0.9117% | 20.96% | 12.43% | 2.097% | 1.107% | 15.19% | 3.002% | 10.08% | 2.995% | 1.425% | 1.887% | 1.622% | 2.2% | 2.099%
43 | Mi1+HL 200520A |  | 5/29/2020 1:27:46 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.060 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 36.58% | 5.13% | 16.91% | 2.176% | 8.957% | 7.202% | 2.366% | 8.439% | 3.639% | 30.7% | 30.7% | 16.08% | 115.9% | 3.617% | 3.522% | 3.525% | 35.64% | 14.51% | 12.79% | 5.322% | 2.468% | 1.624% | 28.5% | 34.22% | 6.288% | 9.081% | 9.718% | 6.345% | 2.307% | 4.022% | 4.152% | 4.659% | 4.45% | S | S | 3.326% | 3.521% | 2.366%
44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.061 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 70.54% | 0.5222% | 12.52% | 2.819% | 1.234% | 1.023% | 0.7956% | 1.741% | 1.247% | 16.66% | 16.66% | 2.95% | 23.22% | 0.3898% | 0.7271% | 1.702% | 53.55% | 8.651% | 11.47% | 1.868% | 0.1985% | 1.17% | 20.4% | 18.13% | 1.198% | 3.365% | 1.457% | 0.7479% | 1.079% | 4.214% | 0.4166% | 3.612% | 0.568% | 0.5439% | S | 0.5644% | 0.9173% | 0.7956%
45 | Rinse |  | 5/29/2020 1:35:06 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.062 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 109% | 6.363% | 28.8% | 32.05% | 11.85% | 6.521% | 1.566% | 6.799% | 29.5% | 7.843% | 7.843% | 54.65% | 30.95% | 218.5% | 30.41% | 0.9284% | 321.1% | 33.91% | 68.66% | 42.83% | 43.9% | 1.023% | 39.24% | 9.177% | 0.8173% | 108% | 14.09% | 50.22% | 0.9997% | 77.13% | 421% | 8.261% | 61.6% | 49.22% | 46.95% | 26.66% | 41.99% | 1.566%
46 | Mi1+HL 200520A 1_99 Spike |  | 5/29/2020 1:38:45 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Spike.063 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.8285% | 4.236% | 0.8357% | 1.088% | 1.476% | 1.508% | 1.542% | 1.396% | 2.096% | 2.112% | 2.112% | 1.36% | 1.085% | 0.8466% | 0.8769% | 0.4564% | 1.955% | 1.167% | 1.013% | 2.057% | 0.1142% | 1.073% | 0.3043% | 1.747% | 2.221% | 0.7764% | 1.943% | 1.235% | 1.961% | 0.7577% | 1.346% | 2.16% | 1.675% | 1.4% | 1.957% | 1.303% | 1.451% | 1.542%
47 | Rinse |  | 5/29/2020 1:42:26 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.064 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 35.28% | 6.718% | 34.02% | 31.26% | 46.82% | 4.828% | 0.9838% | 6.142% | 32.64% | 19.54% | 19.54% | 33.75% | 21.99% | 40.17% | 423.4% | 0.6667% | 203% | 27.58% | 42.87% | 33.09% | 25.27% | 1.527% | 22.75% | 21.78% | 0.4236% | 33.99% | 15.91% | 31.27% | 1.251% | 26.38% | 42.66% | 10.17% | 34.6% | 32.45% | 40.1% | 22.82% | 43.66% | 0.9838%

## Unfactored Concentrations
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (mg/L) | Mg 24 (ug/L) | In 115 (ug/L) | U 238 (mg/L) | CeO 156 (mg/L) | Ce 140 (IS) | Ce++ 70 (mg/L) | Bkgd 220 (ug/L)
1 | Daily Performance Check |  | 5/29/2020 10:21:28 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.018 | C:\NexIONData\Method\Daily Performance.mth
2 | Daily Performance Check |  | 5/29/2020 10:23:57 AM |  | C:\NexIONData\DataSet\052920\Daily Performance Check.019 | C:\NexIONData\Method\Daily Performance.mth
| Sample Id | R | Acquisition Time | QC Status | Dataset File | Method File | Be 9 (ug/L) | B 10 (ug/L) | Al 27 (ug/L) | Ti 47 (ug/L) | V-1 51 (ug/L) | V 51 (ug/L) | Sc 45 (IS) | Cr 52 (ug/L) | Mn 55 (ug/L) | Fe-1 57 (mg/L) | Fe-2 57 (ug/L) | Co 59 (ug/L) | Ni 60 (ug/L) | Cu 65 (ug/L) | Zn 66 (ug/L) | Ge 72 (IS) | As 75 (ug/L) | As-1 75 (ug/L) | Se 82 (ug/L) | Sr 87 (ug/L) | Mo 98 (ug/L) | Rh 103 (IS) | Ag 107 (ug/L) | Cd 111 (ug/L) | In 115 (IS) | Sb 121 (ug/L) | Sn 118 (ug/L) | Ba 137 (ug/L) | Tb 159 (IS) | Tl 205 (ug/L) | Pb 208 (ug/L) | Th 232 (ug/L) | U 238 (ug/L) | Na 23 (ug/L) | Mg 24 (ug/L) | K 39 (ug/L) | Ca 44 (ug/L) | Sc-1 45 (IS) | Cl 35 (ug/L)
3 | Rinse |  | 5/29/2020 10:31:04 AM |  | C:\NexIONData\DataSet\052920\Rinse.020 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
4 | Rinse |  | 5/29/2020 10:37:14 AM |  | C:\NexIONData\DataSet\052920\Rinse.021 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
5 | Rinse |  | 5/29/2020 10:41:20 AM |  | C:\NexIONData\DataSet\052920\Rinse.022 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
6 | Rinse |  | 5/29/2020 10:44:59 AM |  | C:\NexIONData\DataSet\052920\Rinse.023 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
7 | Standard Blank |  | 5/29/2020 10:48:37 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.024 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
8 | Standard Blank |  | 5/29/2020 10:52:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.025 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
9 | 20 ppm Ca |  | 5/29/2020 10:55:53 AM |  | C:\NexIONData\DataSet\052920\20 ppm Ca .026 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
10 | 20 ppm Cl |  | 5/29/2020 10:59:32 AM |  | C:\NexIONData\DataSet\052920\20 ppm Cl.027 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
11 | Rinse |  | 5/29/2020 11:03:12 AM |  | C:\NexIONData\DataSet\052920\Rinse.028 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
12 | Rinse |  | 5/29/2020 11:06:50 AM |  | C:\NexIONData\DataSet\052920\Rinse.029 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
13 | Standard Blank |  | 5/29/2020 11:26:15 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.030 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
14 | Standard Blank |  | 5/29/2020 11:29:53 AM |  | C:\NexIONData\DataSet\052920\Standard Blank.031 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
15 | std blank |  | 5/29/2020 11:33:31 AM |  | C:\NexIONData\DataSet\052920\std blank.032 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth
16 | 1ppb |  | 5/29/2020 11:37:10 AM |  | C:\NexIONData\DataSet\052920\1ppb.033 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 1 | 1 | 1 | 1 | 1 | 1 |  | 1 | 1 | 1 |  | 1 | 1 | 1 | 1 |  | 1 | 1 | 1 | 1 | 1 |  | 1 | 1 |  | 1 | 1 | 1 |  | 1 | 1 | 1 | 1
17 | 10ppb |  | 5/29/2020 11:40:49 AM |  | C:\NexIONData\DataSet\052920\10ppb.034 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 9.9996 | 10.1197 | 9.9776 | 10.0007 | 10.0004 | 10.0006 |  | 10.0014 | 9.9986 | 10.0544 |  | 9.9992 | 9.9984 | 9.9926 | 9.9825 |  | 10.0025 | 10.0018 | 10.0012 | 10.0002 | 10.0003 |  | 10.0157 | 10.0014 |  | 10.0007 | 10.0029 | 9.9979 |  | 10.0003 | 10.0003 | 10.0328 | 9.9993
18 | 50ppb |  | 5/29/2020 11:44:29 AM |  | C:\NexIONData\DataSet\052920\50ppb.035 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 49.8948 | 50.1447 | 49.8679 | 49.9272 | 49.9533 | 49.9467 |  | 49.9145 | 49.9483 | 50.121 |  | 49.9175 | 49.9474 | 49.9142 | 49.9968 |  | 49.859 | 49.8947 | 49.9199 | 49.959 | 49.9069 |  | 49.9164 | 49.8814 |  | 49.8319 | 49.8709 | 49.8879 |  | 49.9023 | 49.9475 | 50.002 | 49.8905
19 | 100ppb |  | 5/29/2020 11:48:09 AM |  | C:\NexIONData\DataSet\052920\100ppb.036 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 98.5071 | 99.2777 | 99.4276 | 99.0529 | 99.5655 | 99.7867 |  | 99.8681 | 99.7109 | 99.7337 |  | 99.8954 | 99.6976 | 99.453 | 99.062 |  | 98.8216 | 98.8709 | 99.2667 | 99.4109 | 101.3564 |  | 100.1313 | 99.1554 |  | 100.1319 | 99.8224 | 99.4085 |  | 99.69 | 99.3085 | 99.543 | 99.3902
20 | 100 PPB |  | 5/29/2020 11:51:50 AM |  | C:\NexIONData\DataSet\052920\100 PPB.037 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.092644 | -1.0974 | 0.18378 | 0.092845 | 0.087536 | 0.063613 |  | 0.074921 | 0.085006 | 104.8994 | 100 | 0.091766 | 0.086912 | 0.15548 | -0.34092 |  | 0.13037 | 0.10151 | 0.12296 | 0.092229 | 0.1082 |  | 0.27404 | 0.093128 |  | 0.090689 | 0.10668 | 0.077583 |  | 0.11064 | 0.087043 | 0.10431 | 0.089793 | 100 | 100 | 100 | 100
21 | 1000 PPB |  | 5/29/2020 11:55:30 AM |  | C:\NexIONData\DataSet\052920\1000 PPB.038 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0080327 | -1.3502 | 0.18766 | 0.033439 | -0.0083206 | -0.019212 |  | 0.015512 | 0.040987 | 1058.2292 | 1000.0864 | 0.034995 | 0.031005 | 0.031955 | -0.41667 |  | 0.0125 | 0.0067046 | 0.021417 | 0.047406 | 0.014668 |  | 0.057938 | 0.019998 |  | 0.0052001 | 0.0099861 | 0.0026905 |  | 0.0084667 | -8.785e-05 | 0.027814 | 0.0049705 | 1000.2123 | 999.8155 | 999.2398 | 1000.6276
22 | 5000 PPB |  | 5/29/2020 11:59:08 AM |  | C:\NexIONData\DataSet\052920\5000 PPB.039 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0032111 | -1.6068 | -0.1507 | 0.14666 | -0.014972 | -0.050407 |  | 0.11088 | 0.2159 | 5288.8638 | 4999.933 | 0.14682 | 0.12912 | 0.8464 | -0.44637 |  | 0.011343 | 0.013206 | 0.026579 | 0.21906 | 0.022553 |  | 0.0274 | 0.024562 |  | 0.0042899 | 0.0074111 | 1.6213 |  | 0.0031406 | 0.00021561 | 0.031411 | 0.0014141 | 4996.849 | 4997.4769 | 4998.7367 | 4996.2896
23 | 10000 PPB |  | 5/29/2020 12:02:47 PM |  | C:\NexIONData\DataSet\052920\10000 PPB.040 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0028459 | -1.5432 | 0.24587 | 0.29911 | -0.013116 | -0.062479 |  | 0.18338 | 0.44032 | 10595.7806 | 10003.4904 | 0.29038 | 0.20358 | 0.42124 | -0.33927 |  | 0.058892 | 0.023951 | 0.030482 | 0.44109 | 0.029416 |  | 0.014084 | 0.029248 |  | 0.0076639 | 0.011791 | 0.039482 |  | 0.0029667 | 0.0051382 | 0.023674 | 0.0017565 | S | 9992.3616 | S | 10019.3082
24 | 20ppb verification std |  | 5/29/2020 12:06:31 PM | Passed | C:\NexIONData\DataSet\052920\20ppb verification std.041 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.5855 | -1.1908 | 20.6169 | 0.037318 | 20.9746 | 21.1279 |  | 21.1545 | 21.0357 | 2031.1531 | 1917.6143 | 20.7356 | 19.9582 | 20.448 | 20.6405 |  | 21.6631 | 21.3196 | 20.787 | 2170.7668 | 19.7069 |  | 20.0917 | 20.2229 |  | 20.3646 | -0.0032876 | 20.6201 |  | 20.1902 | 20.1736 | 20.1351 | 20.5208 | 1986.5761 | 2018.6356 | 2003.7628 | 3236.5155
25 | Ca+B+Sn+Ti check |  | 5/29/2020 12:10:25 PM | Passed | C:\NexIONData\DataSet\052920\Ca+B+Sn+Ti check.042 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.016314 | 19.7907 | 0.054748 | 21.3486 | -0.0086526 | 0.0085047 |  | 0.029981 | 0.0096204 | 1.2049 | 1.1375 | 0.012408 | 0.11415 | 0.022066 | -0.46537 |  | 0.046883 | 0.00048143 | 0.053774 | 1.2355 | 0.017939 |  | 0.099097 | 0.041625 |  | 0.014683 | 21.2194 | 0.013705 |  | 0.016935 | 0.0069514 | 0.039531 | 0.011865 | 0.85384 | 1.25 | 4.7883 | 2098.2652
26 | Verification Check |  | 5/29/2020 12:14:49 PM | Passed | C:\NexIONData\DataSet\052920\Verification Check.043 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.8554 | -1.3116 | 20.3482 | 0.038325 | 20.9218 | 20.7455 |  | 19.6747 | 20.9294 | 2043.205 | 1928.9926 | 20.7501 | 19.8608 | 20.3473 | 20.5691 |  | 21.4719 | 21.2548 | 20.7137 | 2158.511 | 19.8591 |  | 20.3027 | 20.1202 |  | 20.2417 | 0.01746 | 20.4361 |  | 20.543 | 20.4903 | 20.6135 | 21.1095 | 1983.2846 | 2027.5513 | 1991.8059 | 3219.7821
27 | Rinse |  | 5/29/2020 12:18:29 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.044 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.016115 | 6.9019 | 5.5637 | 0.081039 | 0.00045382 | 0.012009 |  | 0.06824 | 0.025475 | 1.0093 | 0.95286 | 0.014923 | 0.020976 | 0.0041816 | -0.018494 |  | 0.038941 | 0.012949 | 0.038302 | 0.99541 | 0.024071 |  | 0.13331 | 0.038055 |  | 0.014122 | 0.033848 | 0.034691 |  | 0.015619 | 0.020672 | 0.12294 | 0.015824 | 4.7441 | 3.3309 | 1.3265 | 19.2931
28 | CRM 1_3 |  | 5/29/2020 12:22:07 PM | Passed | C:\NexIONData\DataSet\052920\CRM 1_3.045 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 5.6746 | -0.82135 | 32.3256 | 0.02744 | 7.4678 | 7.5102 |  | 5.2232 | 10.1262 | 23.4619 | 22.1504 | 6.3331 | 15.1205 | 5.1957 | 20.3411 |  | 21.6361 | 22.6704 | 3.0283 | 63.6926 | 24.6748 |  | 0.44313 | 2.7524 |  | 2.6125 | 0.00088994 | 12.5728 |  | 2.5821 | 10.4265 | 0.1203 | 2.6568 | 1530.5736 | 2307.6853 | 627.9841 | 8687.4921
29 | CRM |  | 5/29/2020 12:25:47 PM | Failed | C:\NexIONData\DataSet\052920\CRM.046 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.1146 | 0.81321 | 127.7002 | 0.087407 | 31.1346 | 31.3357 |  | 21.5262 | 43.8148 | 94.7094 | 89.4153 | 26.9746 | 58.9908 | 19.8002 | 69.0159 |  | 77.6086 | 81.8135 | 9.9844 | 291.1301 | 113.2625 |  | 1.8741 | 9.6859 |  | 9.7034 | 0.016441 | 48.8205 |  | 9.6549 | 38.519 | 0.19583 | 10.3229 | 6038.2467 | 9062.1223 | 2483.0973 | 34263.7294
30 | Rinse |  | 5/29/2020 12:29:27 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.047 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.018509 | 9.2175 | 5.5777 | 0.074712 | 0.022659 | 0.053854 |  | 0.1888 | 0.038762 | 2.058 | 1.943 | 0.020523 | 0.044285 | 0.0021365 | -0.0057285 |  | 0.052238 | 0.045055 | 0.029223 | 0.39518 | 0.087656 |  | 0.06901 | 0.034212 |  | 0.011405 | 0.029411 | 0.047446 |  | 0.011898 | 0.031862 | 0.012541 | 0.012437 | 5.1711 | 7.4185 | 1.9983 | 31.8137
31 | B, Ca, Ti, Sn check |  | 5/29/2020 12:33:05 PM | Passed | C:\NexIONData\DataSet\052920\B, Ca, Ti, Sn check.048 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0030967 | 20.9473 | -0.071951 | 20.6693 | 0.017237 | -0.0078191 |  | 0.039493 | -0.0035839 | -0.1036 | -0.097813 | 0.00096688 | 0.11322 | 0.011896 | -0.39686 |  | 0.01874 | -0.0057445 | 0.0067846 | 0.095176 | 0.0075056 |  | 0.014904 | 0.02021 |  | 0.0014105 | 20.6319 | -0.0025724 |  | 0.0018791 | -0.0050179 | 0.00090619 | 0.00071886 | -2.3402 | -0.12086 | 1.4803 | 2041.3131
32 | Rinse |  | 5/29/2020 12:36:45 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.049 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.007457 | 8.6059 | 5.4038 | 0.08031 | 0.014134 | 0.022311 |  | 0.15355 | 0.014143 | 1.3889 | 1.3113 | 0.0055704 | 0.013108 | -0.013076 | 0.16611 |  | -0.018022 | 0.0044273 | 0.024107 | 0.23559 | 0.013906 |  | 0.01189 | 0.020397 |  | 0.0042948 | 0.054777 | 0.025252 |  | 0.0043557 | 0.009237 | 0.017998 | 0.0054698 | 1.6643 | 2.5294 | 0.076399 | 13.7238
33 | 2 ppb verification check |  | 5/29/2020 12:40:23 PM | Passed | C:\NexIONData\DataSet\052920\2 ppb verification check.050 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 2.2389 | -0.43588 | 2.2541 | 0.0071198 | 2.0282 | 2.0105 |  | 2.1425 | 2.1395 | 202.7299 | 191.3976 | 2.0937 | 2.1016 | 2.1077 | 1.7147 |  | 2.1937 | 2.1708 | 2.0846 | 221.6459 | 2.0114 |  | 1.9023 | 2.0917 |  | 2.0801 | 0.00079256 | 2.0814 |  | 2.0962 | 2.1141 | 1.6022 | 2.1489 | 202.2203 | 208.4265 | 202.8817 | 282.9178
34 | Rinse |  | 5/29/2020 12:44:04 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.051 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.027021 | 8.7792 | 6.6222 | 0.097576 | 0.029877 | 0.057252 |  | 0.20471 | 0.04218 | 3.9116 | 3.6929 | 0.027076 | 0.042645 | 0.040673 | 0.46591 |  | 0.033225 | 0.02734 | 0.055482 | 0.82437 | 0.044817 |  | 0.11521 | 0.039798 |  | 0.022995 | 0.069884 | 0.057125 |  | 0.033942 | 0.059773 | 0.15224 | 0.027839 | 11.8186 | 5.6445 | 4.0276 | 18.8611
35 | Sample blank |  | 5/29/2020 12:47:44 PM | Passed | C:\NexIONData\DataSet\052920\Sample blank.052 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0024318 | -0.75071 | -0.035974 | 0.0018298 | 0.00031077 | -0.010983 |  | 0.014561 | -0.0056516 | -1.0047 | -0.94855 | 0.00040031 | 0.013023 | 0.052144 | -0.48184 |  | 0.026215 | -0.0022449 | -0.0059967 | 0.016713 | 0.0021916 |  | 0.018315 | 0.011875 |  | -3.107e-05 | -0.0032776 | -0.0018151 |  | 0.0011429 | -0.0054567 | 0.0020897 | -0.00011626 | -2.3325 | -0.038933 | 0.44123 | -1.1161
36 | Rinse |  | 5/29/2020 12:51:25 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.053 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0060255 | 7.8422 | 4.8053 | 0.062316 | 0.01206 | 0.025001 |  | 0.15053 | 0.013137 | 0.76034 | 0.71783 | 0.0058093 | 0.011614 | -0.0079795 | -0.16865 |  | -0.033105 | 0.0046332 | -0.0068365 | 0.24756 | 0.010539 |  | 0.012972 | 0.017413 |  | 0.0037897 | 0.028106 | 0.019094 |  | 0.004773 | 0.0084146 | 0.032582 | 0.0060365 | 1.6849 | 2.3528 | -0.81992 | 9.7795
37 | Spike Blank |  | 5/29/2020 12:55:03 PM | Passed | C:\NexIONData\DataSet\052920\Spike Blank.054 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.9983 | 19.8356 | 23.2429 | 20.1556 | 20.8312 | 20.8004 |  | 20.3543 | 20.9857 | 2100.9944 | 1983.5516 | 21.011 | 19.982 | 20.1821 | 20.7961 |  | 21.2172 | 21.1563 | 20.6216 | 20.3563 | 19.6048 |  | 19.8577 | 19.9855 |  | 19.9616 | 19.6221 | 19.8302 |  | 20.5063 | 20.5719 | 20.5922 | 21.153 | 2067.6056 | 2071.6541 | 2019.2891 | 2033.2241
38 | Rinse |  | 5/29/2020 12:58:43 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.055 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.069936 | 8.1065 | 7.4216 | 0.15039 | 0.06459 | 0.079353 |  | 0.16541 | 0.098085 | 8.047 | 7.5971 | 0.074252 | 0.10577 | 0.12906 | 1.5166 |  | 0.075653 | 0.078809 | 0.084906 | 1.8167 | 0.12102 |  | 0.24199 | 0.084189 |  | 0.066338 | 0.16523 | 0.12496 |  | 0.081803 | 0.15701 | 0.31009 | 0.078544 | 25.8421 | 11.5379 | 8.924 | 29.6849
39 | Rinse |  | 5/29/2020 1:02:22 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.056 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0079963 | 7.6462 | 4.7438 | 0.069715 | 4.6526e-05 | 0.012585 |  | 0.10928 | 0.01437 | 0.9795 | 0.92474 | 0.0067061 | 0.013335 | -0.010178 | -0.14894 |  | 0.031653 | 0.0049851 | 0.039991 | 0.22933 | 0.011732 |  | 0.036664 | 0.018662 |  | 0.0060012 | 0.025869 | 0.018807 |  | 0.0062753 | 0.0075991 | 0.054023 | 0.0073409 | 1.5805 | 2.4719 | 0.13515 | 9.0263
40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99.057 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.001809 | -0.14486 | 1.0063 | 0.041753 | -0.0045446 | -0.0021126 |  | 0.022854 | 0.076922 | -0.84139 | -0.79436 | -0.00017029 | -0.0051622 | 0.1023 | -0.07516 |  | -0.015515 | -0.00035806 | 0.012833 | 1.6297 | 0.016463 |  | -0.0023737 | 0.0071866 |  | -0.00050041 | -0.0068597 | 0.69447 |  | 0.00077014 | -0.0029096 | -0.0011939 | 0.036008 | 99.7338 | 407.4063 | 41.2435 | 642.5489
41 | Mi1+HL 200520A 1_99 Duplicate |  | 5/29/2020 1:20:27 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Duplicate.058 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0015333 | 0.15586 | 1.9871 | 0.044144 | 0.00015346 | 0.0039704 |  | 0.035155 | 0.10068 | 0.14742 | 0.13918 | 0.0003706 | 0.0060183 | 0.12585 | 0.41786 |  | 0.0015342 | -0.00096333 | 0.003066 | 1.6587 | 0.018624 |  | 0.027221 | 0.0091401 |  | 0.00068314 | -0.0051492 | 0.69348 |  | 0.00082317 | 0.00049944 | 0.0075813 | 0.037958 | 100.5661 | 414.9264 | 41.6138 | 648.1649
42 | Mi1+HL 200520A 1_9 |  | 5/29/2020 1:24:06 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_9.059 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0021764 | 3.3203 | 2.2023 | 0.143 | 0.015553 | 0.022699 |  | 0.15756 | 0.83937 | 4.6983 | 4.4356 | -0.00055162 | 0.069033 | 1.2162 | 3.6768 |  | 0.054361 | 0.0036635 | 0.062381 | 16.0548 | 0.17343 |  | 0.049559 | 0.0093538 |  | 0.0024667 | 0.0032788 | 6.2793 |  | 0.0032975 | 0.044785 | 0.020842 | 0.36378 | 1004.3906 | 4000.0596 | 402.0685 | 6223.9278
43 | Mi1+HL 200520A |  | 5/29/2020 1:27:46 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.060 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0024714 | 31.5492 | 3.4521 | 0.99786 | 0.10353 | 0.20005 |  | 0.93224 | 8.6508 | 19.1725 | 18.1008 | -0.038949 | 0.09079 | 12.0114 | 39.5073 |  | 0.15112 | 0.053586 | 0.23547 | 188.1734 | 1.9039 |  | 0.048257 | 0.014466 |  | 0.017029 | 0.058216 | 64.6174 |  | 0.014034 | 0.46066 | 0.031411 | 3.8266 | S | S | 4012.1251 | 62391.0665
44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | Failed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A.061 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.00075765 | 30.5901 | 2.8611 | 0.99045 | 0.12841 | 0.18842 |  | 0.93133 | 8.5006 | 20.31 | 19.1747 | -0.036111 | 0.14547 | 11.9127 | 38.8046 |  | 0.098502 | 0.056937 | 0.30557 | 186.6863 | 1.8445 |  | 0.028913 | 0.010731 |  | 0.015438 | 0.052147 | 61.886 |  | 0.013019 | 0.44856 | 0.018073 | 3.6999 | 9672.4492 | S | 3909.0353 | 60836.6124
45 | Rinse |  | 5/29/2020 1:35:06 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.062 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.0026397 | 6.9305 | 4.0785 | 0.055445 | 0.019447 | 0.052996 |  | 0.29508 | 0.015642 | 5.4017 | 5.0998 | 0.0031164 | 0.0096037 | 0.0034768 | -0.26135 |  | 0.019633 | 0.0078027 | 0.043284 | 0.21797 | 0.0068062 |  | 0.0070323 | 0.01541 |  | 0.0013923 | 0.014933 | 0.042504 |  | 0.0018666 | 0.0011744 | 0.029751 | 0.0051989 | 5.1101 | 22.4197 | 2.0488 | 40.0434
46 | Mi1+HL 200520A 1_99 Spike |  | 5/29/2020 1:38:45 PM | Passed | C:\NexIONData\DataSet\052920\Mi1+HL 200520A 1_99 Spike.063 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 20.0315 | 18.4155 | 22.337 | 19.8196 | 20.6693 | 20.5117 |  | 19.6241 | 20.579 | 2072.224 | 1956.3894 | 20.7167 | 19.9348 | 20.5114 | 20.9299 |  | 21.1202 | 21.0252 | 20.5165 | 21.272 | 19.4396 |  | 19.8403 | 19.9352 |  | 19.8953 | 19.5138 | 21.5703 |  | 20.1629 | 20.3751 | 20.1956 | 20.8434 | 2110.6696 | 2400.5851 | 2043.9379 | 2579.7329
47 | Rinse |  | 5/29/2020 1:42:26 PM | Passed | C:\NexIONData\DataSet\052920\Rinse.064 | C:\NexIONData_icpms\Method\HorizonLabTMW (no Hg).mth | 0.01905 | 6.8231 | 5.4822 | 0.081263 | 0.017977 | 0.044549 |  | 0.19496 | 0.029867 | 3.7127 | 3.5052 | 0.017671 | 0.025633 | 0.021544 | -0.034988 |  | 0.011468 | 0.022209 | 0.038445 | 0.35304 | 0.031506 |  | 0.12793 | 0.027561 |  | 0.016299 | 0.051547 | 0.050375 |  | 0.020176 | 0.027256 | 0.15286 | 0.020104 | 6.1973 | 10.7785 | 2.8549 | 21.4814

## Internal Standards
Row Index Number | Sc 45 | Ge 72 | Rh 103 | In 115 | Tb 159 | Sc-1 45
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15 | 100% | 100% | 100% | 100% | 100% | 100%
16 | 99.74% | 99.14% | 99.95% | 98.62% | 100.9% | 99.74%
17 | 100% | 99.55% | 99.33% | 98.97% | 101.6% | 100%
18 | 103.3% | 102% | 102.2% | 104.1% | 103.5% | 103.3%
19 | 105.8% | 102.6% | 102% | 105.5% | 107% | 105.8%
20 | 100.2% | 96.18% | 96.7% | 100.9% | 99.43% | 100.2%
21 | 105.4% | 100.1% | 99.92% | 104.6% | 102.3% | 105.4%
22 | 110.5% | 100.9% | 98.23% | 107.4% | 103.5% | 110.5%
23 | 112.3% | 101.2% | 97.55% | 105.8% | 103.9% | 112.3%
24 | 105.9% | 100.3% | 99.6% | 102.9% | 103% | 105.9%
25 | 101% | 97.87% | 97.13% | 100.2% | 98.88% | 101%
26 | 105.2% | 100.4% | 97.58% | 103.5% | 100.3% | 105.2%
27 | 106.3% | 101.3% | 104% | 106.8% | 109.3% | 106.3%
28 | 102.3% | 97.78% | 95.99% | 100.2% | 96.65% | 102.3%
29 | 109.1% | 96.37% | 95.59% | 102.6% | 101.1% | 109.1%
30 | 105.2% | 99.68% | 100.9% | 103.5% | 103.1% | 105.2%
31 | 102% | 97.76% | 96.71% | 100.9% | 99.43% | 102%
32 | 104.2% | 97.83% | 99.21% | 102.4% | 101.6% | 104.2%
33 | 102% | 97.24% | 97.29% | 100.6% | 98.54% | 102%
34 | 102.1% | 96.26% | 98.36% | 101.6% | 101.3% | 102.1%
35 | 102% | 97.92% | 97.42% | 100.7% | 99.44% | 102%
36 | 105.3% | 99.1% | 100.8% | 104.4% | 104.5% | 105.3%
37 | 107% | 101.5% | 100.2% | 105.1% | 101.4% | 107%
38 | 105.7% | 99.47% | 100.9% | 105% | 105% | 105.7%
39 | 105% | 98.38% | 101.1% | 105.5% | 104.1% | 105%
40 | 101.4% | 97.9% | 95.99% | 101.4% | 98.58% | 101.4%
41 | 101.9% | 99.15% | 97.65% | 103.6% | 100.1% | 101.9%
42 | 109.7% | 101.2% | 99.45% | 107% | 104.1% | 109.7%
43 | 110% | 92.09% | 88.19% | 96.6% | 96.2% | 110%
44 | 111.8% | 92.8% | 90.11% | 99.83% | 97.15% | 111.8%
45 | 104.1% | 96.77% | 99.91% | 102.2% | 102.7% | 104.1%
46 | 103.8% | 98.53% | 96.33% | 102% | 100.1% | 103.8%
47 | 102.4% | 96.38% | 99.82% | 102.8% | 103.5% | 102.4%

## QC
QC Standards
QC STD 1
|  | Sample Id | R | Acquisition Time | Be 9 | B 10 | Al 27 | Ti 47 | V-1 51 | V 51 | Sc 45 (IS) | Cr 52 | Mn 55 | Fe-1 57 | Fe-2 57 | Co 59 | Ni 60 | Cu 65 | Zn 66 | Ge 72 (IS) | As 75 | As-1 75 | Se 82 | Sr 87 | Mo 98 | Rh 103 (IS) | Ag 107 | Cd 111 | In 115 (IS) | Sb 121 | Sn 118 | Ba 137 | Tb 159 (IS) | Tl 205 | Pb 208 | Th 232 | U 238 | Na 23 | Mg 24 | K 39 | Ca 44 | Sc-1 45 (IS) | Cl 35
| 24 | 20ppb verification std |  | 5/29/2020 12:06:31 PM | 20.5855 | -1.1908 | 20.6169 | 0.037318 | 20.9746 | 21.1279 |  | 21.1545 | 21.0357 | 2031.1531 | 1917.6143 | 20.7356 | 19.9582 | 20.448 | 20.6405 |  | 21.6631 | 21.3196 | 20.787 | 2170.7668 | 19.7069 |  | 20.0917 | 20.2229 |  | 20.3646 | -0.0032876 | 20.6201 |  | 20.1902 | 20.1736 | 20.1351 | 20.5208 | 1986.5761 | 2018.6356 | 2003.7628 | 3236.5155
|  | RSD |  |  | 1.992% | 8.032% | 4.631% | 28.74% | 0.9043% | 1.791% |  | 4.532% | 1.203% | 0.8687% | 0.8687% | 0.6243% | 0.4491% | 0.3594% | 0.2453% |  | 1.28% | 0.9152% | 1.435% | 0.4224% | 0.7303% |  | 1.78% | 1.278% |  | 1.481% | 41.18% | 0.7871% |  | 0.8898% | 0.9147% | 1.798% | 0.6059% | 1.36% | 1.4% | 0.8385% | 0.5557%
|  | % Recovery |  |  | 102.9% |  | 103.1% |  | 104.9% |  | 105.9% | 105.8% | 105.2% | 101.6% | 95.88% | 103.7% | 99.79% | 102.2% | 103.2% |  |  | 106.6% | 103.9% | 108.5% | 98.53% | 99.6% | 100.5% | 101.1% | 102.9% | 101.8% |  | 103.1% | 103% | 101% | 100.9% | 100.7% | 102.6% | 99.33% | 100.9% | 100.2% |  | 105.9%
QC STD 3
|  | Sample Id | R | Acquisition Time | Be 9 | B 10 | Al 27 | Ti 47 | V-1 51 | V 51 | Sc 45 (IS) | Cr 52 | Mn 55 | Fe-1 57 | Fe-2 57 | Co 59 | Ni 60 | Cu 65 | Zn 66 | Ge 72 (IS) | As 75 | As-1 75 | Se 82 | Sr 87 | Mo 98 | Rh 103 (IS) | Ag 107 | Cd 111 | In 115 (IS) | Sb 121 | Sn 118 | Ba 137 | Tb 159 (IS) | Tl 205 | Pb 208 | Th 232 | U 238 | Na 23 | Mg 24 | K 39 | Ca 44 | Sc-1 45 (IS) | Cl 35
| 25 | Ca+B+Sn+Ti check |  | 5/29/2020 12:10:25 PM | 0.016314 | 19.7907 | 0.054748 | 21.3486 | -0.0086526 | 0.0085047 |  | 0.029981 | 0.0096204 | 1.2049 | 1.1375 | 0.012408 | 0.11415 | 0.022066 | -0.46537 |  | 0.046883 | 0.00048143 | 0.053774 | 1.2355 | 0.017939 |  | 0.099097 | 0.041625 |  | 0.014683 | 21.2194 | 0.013705 |  | 0.016935 | 0.0069514 | 0.039531 | 0.011865 | 0.85384 | 1.25 | 4.7883 | 2098.2652
|  | RSD |  |  | 29.03% | 4.49% | 493.7% | 3.686% | 34.18% | 118.1% |  | 158% | 25.35% | 57.61% | 57.61% | 16.52% | 8.359% | 9.35% | 3.638% |  | 84.92% | 613.2% | 53.38% | 11.33% | 7.978% |  | 19.27% | 11.99% |  | 15.66% | 2.892% | 30.69% |  | 3.957% | 13.69% | 3.629% | 10.02% | 6.082% | 26.2% | 31.88% | 2.951%
|  | % Recovery |  |  |  | 98.95% |  | 106.7% |  |  | 101% | [14 empty] | 97.13% |  |  | 100.2% |  | 106.1% |  | 98.88% | [7 empty] | 104.9% | 101%
Samples, Blanks & Calibration Standards
|  | Sample Id | R | Acquisition Time | Be 9 | B 10 | Al 27 | Ti 47 | V-1 51 | V 51 | Sc 45 (IS) | Cr 52 | Mn 55 | Fe-1 57 | Fe-2 57 | Co 59 | Ni 60 | Cu 65 | Zn 66 | Ge 72 (IS) | As 75 | As-1 75 | Se 82 | Sr 87 | Mo 98 | Rh 103 (IS) | Ag 107 | Cd 111 | In 115 (IS) | Sb 121 | Sn 118 | Ba 137 | Tb 159 (IS) | Tl 205 | Pb 208 | Th 232 | U 238 | Na 23 | Mg 24 | K 39 | Ca 44 | Sc-1 45 (IS) | Cl 35
| 29 | CRM |  | 5/29/2020 12:25:47 PM | 20.1146 | 0.81321 | 127.7002 | 0.087407 | 31.1346 | 31.3357 |  | 21.5262 | 43.8148 | 94.7094 | 89.4153 | 26.9746 | 58.9908 | 19.8002 | 69.0159 |  | 77.6086 | 81.8135 | 9.9844 | 291.1301 | 113.2625 |  | 1.8741 | 9.6859 |  | 9.7034 | 0.016441 | 48.8205 |  | 9.6549 | 38.519 | 0.19583 | 10.3229 | 6038.2467 | 9062.1223 | 2483.0973 | 34263.7294
|  | IS % Recovery | [8 empty] | 109.1% | [14 empty] | 95.59% |  |  | 102.6% |  |  |  | 101.1% | [8 empty] | 109.1%
|  | Sample RSD |  |  | 2.566% | 85.28% | 2.245% | 11.46% | 1.879% | 2.747% |  | 6.265% | 0.8802% | 2.283% | 2.283% | 0.4687% | 0.8986% | 0.9422% | 0.6217% |  | 0.474% | 0.4934% | 1.016% | 0.713% | 2.265% |  | 5.422% | 1.334% |  | 1.532% | 17.95% | 1.379% |  | 2.53% | 2.671% | 7.66% | 3.363% | 2.463% | 2.612% | 1.712% | 1.141%
|  | Sample Concentration SD |  |  | 0.51611 | 0.69353 | 2.8665 | 0.010021 | 0.58488 | 0.86073 |  | 1.3487 | 0.38564 | 2.1624 | 2.0415 | 0.12644 | 0.53008 | 0.18656 | 0.4291 |  | 0.3679 | 0.40366 | 0.10145 | 2.0757 | 2.5655 |  | 0.10162 | 0.12925 |  | 0.14867 | 0.0029506 | 0.67307 |  | 0.24422 | 1.0287 | 0.015001 | 0.34717 | 148.72 | 236.6962 | 42.523 | 391.0962
| 40 | Mi1+HL 200520A 1_99 |  | 5/29/2020 1:16:47 PM | 0.1809 | -14.4855 | 100.632 | 4.1753 | -0.45446 | -0.21126 |  | 2.2854 | 7.6922 | -84.1389 | -79.4357 | -0.017029 | -0.51622 | 10.2296 | -7.516 |  | -1.5515 | -0.035806 | 1.2833 | 162.9652 | 1.6463 |  | -0.23737 | 0.71866 |  | -0.050041 | -0.68597 | 69.4473 |  | 0.077014 | -0.29096 | -0.11939 | 3.6008 | 9973.3778 | 40740.6268 | 4124.3476 | 64254.888
|  | IS % Recovery | [8 empty] | 101.4% | [14 empty] | 95.99% |  |  | 101.4% |  |  |  | 98.58% | [8 empty] | 101.4%
|  | Sample RSD |  |  | 59.12% | 197.3% | 28.77% | 36.71% | 63.49% | 337% |  | 64.36% | 3.323% | 60.62% | 60.62% | 71.41% | 37.41% | 7.384% | 26.88% |  | 111.6% | 390.3% | 55.17% | 2.282% | 3.178% |  | 13.05% | 49.18% |  | 47.9% | 2.376% | 2.49% |  | 29.93% | 27% | 14.06% | 1.627% | 1.543% | 1.309% | 3.039% | 1.211%
|  | Sample Concentration SD |  |  | 0.10695 | 28.5743 | 28.9566 | 1.5326 | 0.28856 | 0.71204 |  | 1.4708 | 0.25565 | 51.0011 | 48.1502 | 0.01216 | 0.1931 | 0.75538 | 2.0204 |  | 1.7323 | 0.13974 | 0.70804 | 3.7182 | 0.052324 |  | 0.030969 | 0.35347 |  | 0.023969 | 0.016301 | 1.7293 |  | 0.023049 | 0.078569 | 0.016781 | 0.058573 | 153.9284 | 533.213 | 125.3276 | 777.8942
| 43 | Mi1+HL 200520A |  | 5/29/2020 1:27:46 PM | 0.0024714 | 31.5492 | 3.4521 | 0.99786 | 0.10353 | 0.20005 |  | 0.93224 | 8.6508 | 19.1725 | 18.1008 | -0.038949 | 0.09079 | 12.0114 | 39.5073 |  | 0.15112 | 0.053586 | 0.23547 | 188.1734 | 1.9039 |  | 0.048257 | 0.014466 |  | 0.017029 | 0.058216 | 64.6174 |  | 0.014034 | 0.46066 | 0.031411 | 3.8266 | S | S | 4012.1251 | 62391.0665
|  | IS % Recovery | [8 empty] | 110% | [14 empty] | 88.19% |  |  | 96.6% |  |  |  | 96.2% | [8 empty] | 110%
|  | Sample RSD |  |  | 36.58% | 5.13% | 16.91% | 2.176% | 8.957% | 7.202% |  | 8.439% | 3.639% | 30.7% | 30.7% | 16.08% | 115.9% | 3.617% | 3.522% |  | 35.64% | 14.51% | 12.79% | 5.322% | 2.468% |  | 28.5% | 34.22% |  | 9.081% | 9.718% | 6.345% |  | 4.022% | 4.152% | 4.659% | 4.45% | S | S | 3.326% | 3.521%
|  | Sample Concentration SD |  |  | 0.00090405 | 1.6184 | 0.58361 | 0.021718 | 0.009273 | 0.014409 |  | 0.078673 | 0.31484 | 5.8867 | 5.5576 | 0.0062639 | 0.10521 | 0.4344 | 1.3915 |  | 0.05385 | 0.0077745 | 0.030115 | 10.0146 | 0.046986 |  | 0.013751 | 0.0049507 |  | 0.0015463 | 0.0056577 | 4.1001 |  | 0.00056444 | 0.019127 | 0.0014635 | 0.17029 | S | S | 133.4367 | 2196.6136
| 44 | Mi1+HL 200520A |  | 5/29/2020 1:31:26 PM | 0.00075765 | 30.5901 | 2.8611 | 0.99045 | 0.12841 | 0.18842 |  | 0.93133 | 8.5006 | 20.31 | 19.1747 | -0.036111 | 0.14547 | 11.9127 | 38.8046 |  | 0.098502 | 0.056937 | 0.30557 | 186.6863 | 1.8445 |  | 0.028913 | 0.010731 |  | 0.015438 | 0.052147 | 61.886 |  | 0.013019 | 0.44856 | 0.018073 | 3.6999 | 9672.4492 | S | 3909.0353 | 60836.6124
|  | IS % Recovery | [8 empty] | 111.8% | [14 empty] | 90.11% |  |  | 99.83% |  |  |  | 97.15% | [8 empty] | 111.8%
|  | Sample RSD |  |  | 70.54% | 0.5222% | 12.52% | 2.819% | 1.234% | 1.023% |  | 1.741% | 1.247% | 16.66% | 16.66% | 2.95% | 23.22% | 0.3898% | 0.7271% |  | 53.55% | 8.651% | 11.47% | 1.868% | 0.1985% |  | 20.4% | 18.13% |  | 3.365% | 1.457% | 0.7479% |  | 4.214% | 0.4166% | 3.612% | 0.568% | 0.5439% | S | 0.5644% | 0.9173%
|  | Sample Concentration SD |  |  | 0.00053447 | 0.15973 | 0.3583 | 0.027918 | 0.001584 | 0.0019274 |  | 0.016211 | 0.10604 | 3.3843 | 3.1951 | 0.0010655 | 0.033774 | 0.046434 | 0.28216 |  | 0.052752 | 0.0049259 | 0.035061 | 3.4879 | 0.0036623 |  | 0.0058984 | 0.0019459 |  | 0.0005195 | 0.00075967 | 0.46285 |  | 0.00054868 | 0.0018688 | 0.00065276 | 0.021015 | 52.6041 | S | 22.0612 | 558.0288